## TEa & SA
ANALYTE | SCHEME | BASIS TEa | EVALUATION LEVEL | UNIT | TEa (%) | SA (%) | TEa | SA | BASE FOR QUANTITATIVE SCORE | LAST TE CHANGE
Acanthamoeba PCR | Acanthamoeba Extern. Mol. Quality Assessment |  | 20 | {Cq}
ACE | ACE | Biological variation | 20 | U/L | 17.91 | 69.4131 | 3.582 | 13.8826 | SDsa1
Acetone | Alcohols/GHB | Based on EMA/FDA | 1 | g/L | 26.5 | 19.4251 | 0.265 | 0.19425 | TEa
Ethanol | Alcohols/GHB | Based on EMA/FDA | 5 | g/L | 26.5 | 12.3024 | 1.325 | 0.61512 | TEa
Ethylene glycol | Alcohols/GHB | Based on EMA/FDA | 1 | g/L | 26.5 | 26.9777 | 0.265 | 0.26978 | SDsa1
GHB | Alcohols/GHB | Based on EMA/FDA | 1000 | mg/L | 26.5 | 28.5136 | 265 | 285.1365 | SDsa1
Isopropanol | Alcohols/GHB | Based on EMA/FDA | 1 | g/L | 26.5 | 13.7754 | 0.265 | 0.13775 | TEa
Methanol | Alcohols/GHB | Based on EMA/FDA | 1 | g/L | 26.5 | 13.8956 | 0.265 | 0.13896 | TEa
Type III allergie IgG AB (Qual.) | Allergy Type III antibodies |  |  | Qualit.
Type III allergy IgG AB | Allergy Type III antibodies |  | 10 | mg/L |  |  |  |  | TE
ASXL1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
BCOR mutation | AML molecular diagnostics |  | 5 | %{VAF}
CBFB-MYH11 | AML molecular diagnostics |  |  | Qualit.
CEBPA bZIP mutation | AML molecular diagnostics |  | 5 | %{VAF}
CEBPA dm | AML molecular diagnostics |  | 5 | %{VAF}
EZH2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
FLT3 ITD | AML molecular diagnostics |  | 2.5 | {mutant_wt_ratio} |  |  |  |  | SDsa1
FLT3 TKD | AML molecular diagnostics |  | 0.25 | {mutant_wt_ratio} |  |  |  |  | SDsa1
GATA2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
IDH1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
IDH2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
KIT mutation | AML molecular diagnostics |  | 5 | %{VAF}
NPM1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
RUNX1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
RUNX1-RUNX1T1 | AML molecular diagnostics |  |  | Qualit.
SF3B1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
SRSF2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
STAG2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
TP53 mutation | AML molecular diagnostics |  | 5 | %{VAF}
U2AF1 mutation | AML molecular diagnostics |  | 5 | %{VAF}
ZRSR2 mutation | AML molecular diagnostics |  | 5 | %{VAF}
Acetaminophen | Analgesic drugs | Based on EMA/FDA | 100 | mg/L | 26.5 | 17.9959 | 26.5 | 17.9959 | TEa
Ibuprofen | Analgesic drugs | Based on EMA/FDA | 50 | mg/L | 26.5 | 26.3979 | 13.25 | 13.1989 | TEa
Salicylic acid | Analgesic drugs | Based on EMA/FDA | 250 | mg/L | 26.5 | 12.857 | 66.25 | 32.1425 | TEa
ANA IIF interfererend met ANCA | ANCA-GBM |  |  | Qualit.
ANCA IIF (ethanol gefixeerd) | ANCA-GBM |  |  | Qualit.
ANCA IIF (formaline ANCA IIF (formaline gefixeerd) | ANCA-GBM |  |  | Qualit.
ANCA IIF conclusie | ANCA-GBM |  | 300 | {titer} |  | 51.9057 |  | 103.8113 | SDsa1
antibodies against GBM | ANCA-GBM |  | 200 | [IU]/mL |  | 52.4651 |  | 52.4651 | SDsa1
antibodies against MPO | ANCA-GBM |  | 100 | [IU]/mL |  | 95.5395 |  | 95.5395 | SDsa1
antibodies against PR3 | ANCA-GBM |  | 25 | [IU]/mL |  | 49.7061 |  | 12.4265 | SDsa1
Sterkte van het patroon | ANCA-GBM |  |  | Qualit.
Amikacin | Antibiotic drugs | Based on EMA/FDA | 25 | mg/L | 26.5 | 10.3263 | 6.625 | 2.5816 | TEa
Gentamicin | Antibiotic drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 21.1386 | 2.65 | 2.1139 | TEa
Methotrexate | Antibiotic drugs | Based on EMA/FDA | 30 | umol/L | 26.5 | 23.0551 | 7.95 | 6.9165 | TEa
Tobramycin | Antibiotic drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 17.2965 | 2.65 | 1.7297 | TEa |  | Difference plot kreatinine: green area TEa interval, blue area area SDsa1 interval.
Vancomycin | Antibiotic drugs | Based on EMA/FDA | 20 | mg/L | 26.5 | 18.6893 | 5.3 | 3.7379 | TEa
Ceftazidim | Antibiotic drugs 2 | Based on EMA/FDA | 80 | mg/L | 26.5 | 31.9657 | 21.2 | 25.5726 | SDsa1
Ciprofloxacine | Antibiotic drugs 2 | Based on EMA/FDA | 0.7 | mg/L | 26.5 | 36.4733 | 0.1855 | 0.25531 | SDsa1
Flucloxacilline | Antibiotic drugs 2 | Based on EMA/FDA | 6 | mg/L | 26.5 | 24.3587 | 1.59 | 1.4615 | TEa
N-acetyl Sulfamethoxazol | Antibiotic drugs 2 | Based on EMA/FDA | 160 | mg/L | 26.5 |  | 42.4 |  | SDsa1
Piperacilline | Antibiotic drugs 2 | Based on EMA/FDA | 3 | mg/L | 26.5 | 32.9099 | 0.795 | 0.9873 | SDsa1
protein-unbound (free) flucloxacillin | Antibiotic drugs 2 | Based on EMA/FDA | 3 | mg/L | 26.5 |  | 0.795 |  | SDsa1
Sulfamethoxazol | Antibiotic drugs 2 | Based on EMA/FDA | 160 | mg/L | 26.5 | 34.5231 | 42.4 | 55.2369 | SDsa1
Trimethoprim | Antibiotic drugs 2 | Based on EMA/FDA | 8 | mg/L | 26.5 | 19.7911 | 2.12 | 1.5833 | TEa
Antibodies against desmogleïne 1 | Antibodies against skin |  |  | {titer}
Antibodies against desmogleïne 3 | Antibodies against skin |  |  | {titer}
Antibodies against NC16A domain BP180 | Antibodies against skin |  |  | {titer}
Basement membrane (BMZ) | Antibodies against skin |  |  | Qualit.
extracellu. surface keratinocytes (ECS) | Antibodies against skin |  |  | Qualit.
Carbamazepine | Antiepileptic drugs | Based on EMA/FDA | 8 | mg/L | 26.5 | 14.6353 | 2.12 | 1.1708 | TEa
Epoxide-cbz | Antiepileptic drugs | Based on EMA/FDA | 1 | mg/L | 26.5 | 21.3395 | 0.265 | 0.21339 | TEa
Ethosuximide | Antiepileptic drugs | Based on EMA/FDA | 70 | mg/L | 26.5 | 20.5861 | 18.55 | 14.4103 | TEa
Lamotrigine | Antiepileptic drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 18.8108 | 2.65 | 1.8811 | TEa
Levetiracetam | Antiepileptic drugs | Based on EMA/FDA | 15 | mg/L | 26.5 | 20.5231 | 3.975 | 3.0785 | TEa
Mono-hydroxy-carbazepine | Antiepileptic drugs | Based on EMA/FDA | 25 | mg/L | 26.5 | 17.9474 | 6.625 | 4.4869 | TEa
Phenobarbital | Antiepileptic drugs | Based on EMA/FDA | 30 | mg/L | 26.5 | 13.654 | 7.95 | 4.0962 | TEa
Phenytoin | Antiepileptic drugs | Based on EMA/FDA | 15 | mg/L | 26.5 | 17.7618 | 3.975 | 2.6643 | TEa
Valproic acid | Antiepileptic drugs | Based on EMA/FDA | 75 | mg/L | 26.5 | 16.2066 | 19.875 | 12.1549 | TEa
Phenytoin | Antiepileptic drugs free concentration | Based on EMA/FDA | 15 | mg/L | 26.5 | 17.7618 | 3.975 | 2.6643 | TEa
Phenytoïn free concentration | Antiepileptic drugs free concentration | Based on EMA/FDA | 1.5 | mg/L | 26.5 | 25.0048 | 0.3975 | 0.37507 | TEa
Valproic acid | Antiepileptic drugs free concentration | Based on EMA/FDA | 75 | mg/L | 26.5 | 16.2066 | 19.875 | 12.1549 | TEa
Valproic acid free concentration | Antiepileptic drugs free concentration | Based on EMA/FDA | 7.5 | mg/L | 26.5 | 15.447 | 1.9875 | 1.1585 | TEa
Fluconazole | Antifungal drugs | Based on EMA/FDA | 15 | mg/L | 26.5 | 28.079 | 3.975 | 4.2119 | SDsa1
Flucytosine | Antifungal drugs | Based on EMA/FDA | 50 | mg/L | 26.5 | 15.4375 | 13.25 | 7.7188 | TEa
Hydroxy-itraconazole | Antifungal drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 25.5207 | 0.53 | 0.51041 | TEa
Isavuconazol | Antifungal drugs | Based on EMA/FDA | 1 | mg/L | 26.5 | 24.5682 | 0.265 | 0.24568 | TEa
Itraconazole | Antifungal drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 20.9547 | 0.53 | 0.41909 | TEa
Posaconazole | Antifungal drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 26.5301 | 0.53 | 0.5306 | SDsa1
Voriconazole | Antifungal drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 22.3952 | 0.53 | 0.4479 | TEa
Amprenavir | Antiretroviral drugs | Based on EMA/FDA | 3 | mg/L | 26.5 | 15.6743 | 0.795 | 0.47023 | TEa
Atazanavir | Antiretroviral drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 24.5244 | 0.53 | 0.49049 | TEa
Darunavir | Antiretroviral drugs | Based on EMA/FDA | 5 | mg/L | 26.5 | 24.6513 | 1.325 | 1.2326 | TEa
Dolutegravir | Antiretroviral drugs | Based on EMA/FDA | 1.5 | mg/L | 26.5 | 28.0441 | 0.3975 | 0.42066 | SDsa1
Efavirenz | Antiretroviral drugs | Based on EMA/FDA | 6 | mg/L | 26.5 | 18.0474 | 1.59 | 1.0828 | TEa
Elvitegravir | Antiretroviral drugs | Based on EMA/FDA | 0.75 | mg/L | 26.5 | 20.3408 | 0.19875 | 0.15256 | TEa
Etravirine | Antiretroviral drugs | Based on EMA/FDA | 0.5 | mg/L | 26.5 | 32.2719 | 0.1325 | 0.16136 | SDsa1
Lopinavir | Antiretroviral drugs | Based on EMA/FDA | 5 | mg/L | 26.5 | 17.6942 | 1.325 | 0.88471 | TEa
Nevirapine | Antiretroviral drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 22.425 | 2.65 | 2.2425 | TEa
Raltegravir | Antiretroviral drugs | Based on EMA/FDA | 4 | mg/L | 26.5 | 24.4442 | 1.06 | 0.97777 | TEa
Rilpivirine | Antiretroviral drugs | Based on EMA/FDA | 0.5 | mg/L | 26.5 | 24.7179 | 0.1325 | 0.12359 | TEa
Saquinavir | Antiretroviral drugs | Based on EMA/FDA | 5 | mg/L | 26.5 | 23.5189 | 1.325 | 1.1759 | TEa
Tipranavir | Antiretroviral drugs | Based on EMA/FDA | 40 | mg/L | 26.5 | 22.1201 | 10.6 | 8.848 | TEa
acetyl-isoniazide | Antituberculosis drugs | Based on EMA/FDA | 3 | mg/L | 26.5 |  | 0.795 |  | TEa
Bedaquiline | Antituberculosis drugs | Based on EMA/FDA | 1.5 | mg/L | 26.5 |  | 0.3975 |  | SDsa1
clofazimine | Antituberculosis drugs | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265 |  | SDsa1
Delamanid | Antituberculosis drugs | Based on EMA/FDA | 0.1 | mg/L | 26.5 |  | 0.0265 |  | SDsa1
Desacetylrifabutin | Antituberculosis drugs | Based on EMA/FDA | 0.06 | mg/L | 26.5 |  | 0.0159 |  | SDsa1
Ethambutol | Antituberculosis drugs | Based on EMA/FDA | 3 | mg/L | 26.5 | 26.7865 | 0.795 | 0.80359 | SDsa1
Isoniazid | Antituberculosis drugs | Based on EMA/FDA | 3 | mg/L | 26.5 | 22.8705 | 0.795 | 0.68612 | TEa
Levofloxacin | Antituberculosis drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 14.9373 | 2.65 | 1.4937 | TEa
Linezolid | Antituberculosis drugs | Based on EMA/FDA | 8 | mg/L | 26.5 | 25.0586 | 2.12 | 2.0047 | TEa
Moxifloxacin | Antituberculosis drugs | Based on EMA/FDA | 4 | mg/L | 26.5 |  | 1.06 |  | SDsa1
Pyrazinamide | Antituberculosis drugs | Based on EMA/FDA | 40 | mg/L | 26.5 | 23.7332 | 10.6 | 9.4933 | TEa
rifabutin | Antituberculosis drugs | Based on EMA/FDA | 0.6 | mg/L | 26.5 |  | 0.159 |  | SDsa1
Rifampicin | Antituberculosis drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 23.712 | 2.65 | 2.3712 | TEa
Aciclovir | Antiviral drugs | Based on EMA/FDA | 5 | mg/L | 26.5 | 16.1264 | 1.325 | 0.80632 | TEa
Ganciclovir | Antiviral drugs | Based on EMA/FDA | 5 | mg/L | 26.5 | 21.151 | 1.325 | 1.0576 | TEa
Ribavirin | Antiviral drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 16.5131 | 0.53 | 0.33026 | TEa
AF-bone | AP-isoenzymes | Biological variation | 45 | % | 14.497 | 39.0222 | 6.5236 | 17.56 | SDsa1 | 2021-01-01
AF-Gal | AP-isoenzymes |  | 0.5 | % |  |  |  |  | SDsa1
AF-intestinal | AP-isoenzymes |  | 2 | % |  |  |  |  | SDsa1
AF-liver 1 | AP-isoenzymes | Biological variation | 44 | % | 10.41 | 35.8292 | 4.5804 | 15.7648 | SDsa1 | 2024-01-25
AF-liver 2 | AP-isoenzymes | Biological variation | 8 | % | 10.41 | 21.2839 | 0.8328 | 1.7027 | SDsa1 | 2024-01-25
AF-Macromoleculaire fraction | AP-isoenzymes |  | 2 | % |  |  |  |  | SDsa1
Other | AP-isoenzymes |  |  | % |  |  |  |  | TEa
Anti-DNaseB | AST / AntiDNaseB serology |  |  | IZS **
AST | AST / AntiDNaseB serology |  |  | IZS **
Other(specify at comments) | AST / AntiDNaseB serology |  |  | IZS **
Wijst op recente streptokokkeninf. | AST / AntiDNaseB serology |  |  | Qualit.
Determination Bacteriology | Bacteriology |  |  | Qualit.
G. Bac. clindamycine (measured) | Bacteriology |  |  | Qualit.
G. Bac. clindamycine (reporting) | Bacteriology |  |  | Qualit.
G. Bac. meropenem (measured) | Bacteriology |  |  | Qualit.
G. Bac. meropenem (reporting) | Bacteriology |  |  | Qualit.
G. Bac. metronidazole (measured) | Bacteriology |  |  | Qualit.
G. Bac. metronidazole (reporting) | Bacteriology |  |  | Qualit.
G. Bac. piperacilline/tazobactam (measured) | Bacteriology |  |  | Qualit.
G. Bac. piperacilline/tazobactam (reporting) | Bacteriology |  |  | Qualit.
G. Camp. Tetracycline (measured) | Bacteriology |  |  | Qualit.
G. Cory. Rifampicin (measured) | Bacteriology |  |  | Qualit.
G. Cory. Rifampicin (report) | Bacteriology |  |  | Qualit.
G. Kleb. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
G. Kluy. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
G. Kluy. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
G. Sta. aur. Mupirocine (measured) | Bacteriology |  |  | Qualit.
G. Sta. aur. Mupirocine (reporting) | Bacteriology |  |  | Qualit.
G. Sta. spp. oxacilline (Measured) | Bacteriology |  |  | Qualit.
G. Sta. spp. oxacilline (Reporting) | Bacteriology |  |  | Qualit.
G. Str. clindamycine (measured) | Bacteriology |  |  | Qualit.
G. Str. clindamycine (Reporting) | Bacteriology |  |  | Qualit.
G. Str. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
G. Str. Co-trimoxazol (Raporting) | Bacteriology |  |  | Qualit.
G. Str. doxycycline (Measured) | Bacteriology |  |  | Qualit.
G. Str. doxycycline (Reporting) | Bacteriology |  |  | Qualit.
G. Str. erythromycine (Measured) | Bacteriology |  |  | Qualit.
G. Str. erythromycine (Reporting) | Bacteriology |  |  | Qualit.
G. Str. Penicilline (Measured) | Bacteriology |  |  | Qualit.
G. Str. Penicilline (Reporting) | Bacteriology |  |  | Qualit.
G. Str. tetracycline (measured) | Bacteriology |  |  | Qualit.
G. Str. tetracycline (reporting) | Bacteriology |  |  | Qualit.
G. Str. Vancomycin (Measured) | Bacteriology |  |  | Qualit.
G. Str. Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
RM. Aci. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Cit. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Cit. ESBL | Bacteriology |  |  | Qualit.
RM. Cit. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. E. faecium High level Gentamicin | Bacteriology |  |  | Qualit.
RM. E. faecium resistance (HLGR) | Bacteriology |  |  | Qualit.
RM. E. faecium VRE | Bacteriology |  |  | Qualit.
RM. E.faecalis High level Gentamicin | Bacteriology |  |  | Qualit.
RM. E.faecalis resistance (HLGR) | Bacteriology |  |  | Qualit.
RM. E.faecalis VRE | Bacteriology |  |  | Qualit.
RM. Entbac. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Entbac. ESBL | Bacteriology |  |  | Qualit.
RM. Entbac. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Esch. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Esch. ESBL | Bacteriology |  |  | Qualit.
RM. Esch. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Haf. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Haf. ESBL | Bacteriology |  |  | Qualit.
RM. Haf. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Kleb. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Kleb. ESBL | Bacteriology |  |  | Qualit.
RM. Kleb. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Kluy. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Kluy. ESBL | Bacteriology |  |  | Qualit.
RM. Kluy. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Lecl. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Lecl. ESBL | Bacteriology |  |  | Qualit.
RM. Lecl. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Morg. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Morg. ESBL | Bacteriology |  |  | Qualit.
RM. Morg. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Prot. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Prot. ESBL | Bacteriology |  |  | Qualit.
RM. Prot. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Prov. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Prov. ESBL | Bacteriology |  |  | Qualit.
RM. Prov. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Pseu. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Salm. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Salm. ESBL | Bacteriology |  |  | Qualit.
RM. Salm. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Ser. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Ser. ESBL | Bacteriology |  |  | Qualit.
RM. Ser. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Shig. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Shig. ESBL | Bacteriology |  |  | Qualit.
RM. Shig. Plasmid AmpC | Bacteriology |  |  | Qualit.
RM. Sta. aur. MLSB inducibale | Bacteriology |  |  | Qualit.
RM. Sta. aur. MRSA | Bacteriology |  |  | Qualit.
RM. Sta spp. MLSB inducible | Bacteriology |  |  | Qualit.
RM. Str. MLSB inducible | Bacteriology |  |  | Qualit.
RM. Str. pne. MLSB inducible | Bacteriology |  |  | Qualit.
RM. Yer. Carbapenemase | Bacteriology |  |  | Qualit.
RM. Yer. ESBL | Bacteriology |  |  | Qualit.
RM. Yer. Plasmid AmpC | Bacteriology |  |  | Qualit.
S. Aci. Amikacin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Amikacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Colistin (Reportin) | Bacteriology |  |  | Qualit.
S. Aci. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Aci. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Aci. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Levofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Levofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Aci. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Aci. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Aci. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Aec. Ampicillin (measured) | Bacteriology |  |  | Qualit.
S. Aec. Ampicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Aec. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Aec. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aec. Nitrofurantoin (measured) | Bacteriology |  |  | Qualit.
S. Aec. Nitrofurantoin (Reporting) | Bacteriology |  |  | Qualit.
S. Aec. Norfloxacin (measured) | Bacteriology |  |  | Qualit.
S. Aec. Norfloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aec. Penicillin (measured) | Bacteriology |  |  | Qualit.
S. Aec. Penicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Aztreonam (measuring) | Bacteriology |  |  | Qualit.
S. Aer. Aztreonam (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Cefepime (measuring) | Bacteriology |  |  | Qualit.
S. Aer. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Aer. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Aer. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Aer. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Aer. Levofloxacine (measuring) | Bacteriology |  |  | Qualit.
S. Aer. Levofloxacine (Reporting) | Bacteriology |  |  | Qualit.
S. Bac. amoxicilline/clavulaanzuur (measured) | Bacteriology |  |  | Qualit.
S. Bac. amoxicilline/clavulaanzuur (reporting) | Bacteriology |  |  | Qualit.
S. Camp. Azithromycin (measured) | Bacteriology |  |  | Qualit.
S. Camp. Azithromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Camp. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Camp. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Camp. Clarithromycin (measured) | Bacteriology |  |  | Qualit.
S. Camp. Clarithromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Camp. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Camp. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Camp. Erythromycin (measured) | Bacteriology |  |  | Qualit.
S. Camp. Erythromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Camp. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Cit. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Cit. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Cit. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Cit. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Cit. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Cit. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Cit. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Cit. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Cit. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Cit. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Clos. Metronidazole (measured) | Bacteriology |  |  | Qualit.
S. Clos. Metronidazole (Reporting) | Bacteriology |  |  | Qualit.
S. Clos. Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. Clos. Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Cory. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Clindamycin (measured) | Bacteriology |  |  | Qualit.
S. Cory. Clindamycin (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Cory. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Penicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Penicilline (measured) | Bacteriology |  |  | Qualit.
S. Cory. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Cory. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Cory. Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. Cory. Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecalis Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. E.faecalis Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecalis Ampicillin (measured) | Bacteriology |  |  | Qualit.
S. E.faecalis Ampicillin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecalis Nitrofurantion (measured) | Bacteriology |  |  | Qualit.
S. E.faecalis Nitrofurantion (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecalis Teicoplanin (measured) | Bacteriology |  |  | Qualit.
S. E.faecalis Teicoplanin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecalis Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. E.faecalis Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecium Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. E.faecium Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecium Ampicillin (measured) | Bacteriology |  |  | Qualit.
S. E.faecium Ampicillin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecium Teicoplanin (measured) | Bacteriology |  |  | Qualit.
S. E.faecium Teicoplanin (Reporting) | Bacteriology |  |  | Qualit.
S. E.faecium Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. E.faecium Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Entbac. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Entbac. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Esch. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Esch. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Esch. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Esch. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Esch. Cefuroxime (Rapportage) | Bacteriology |  |  | Qualit.
S. Esch. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Esch. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Esch. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Esch. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Nitrofurantoin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Nitrofurantoin (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Esch. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Esch. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Esch. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ampicillin (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ampicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. inf. Rifampicin (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Rifampicin (Rapportage) | Bacteriology |  |  | Qualit.
S. Hae. inf. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Hae. inf. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Ampicillin (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Ampicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Hae. par. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Hae. par. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Haf. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Haf. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Haf. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Haf. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Haf. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Haf. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Haf. Meropenem (Rapportage) | Bacteriology |  |  | Qualit.
S. Haf. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Haf. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Haf. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Haf. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Kleb. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Kleb. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Kluy. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Kluy. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Cefuroxime (Raporting) | Bacteriology |  |  | Qualit.
S. Lecl. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Lecl. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Lecl. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. amoxicilline (Measured) | Bacteriology |  |  | Qualit.
S. Lis. amoxicilline (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. ampicilline (Measured) | Bacteriology |  |  | Qualit.
S. Lis. ampicilline (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. erythromycine (Measured) | Bacteriology |  |  | Qualit.
S. Lis. erythromycine (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. meropenem (Measured) | Bacteriology |  |  | Qualit.
S. Lis. meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. penicilline (Measured) | Bacteriology |  |  | Qualit.
S. Lis. penicilline (Reporting) | Bacteriology |  |  | Qualit.
S. Lis. trimethoprim-sulfamethoxazole (measured) | Bacteriology |  |  | Qualit.
S. Lis. trimethoprim-sulfamethoxazole (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Morg. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Morg. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Morg. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Morg. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Morg. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Morg. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Morg. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Morg. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Morg. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Morg. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. gon. Azithromycin (measured) | Bacteriology |  |  | Qualit.
S. Neis. gon. Azithromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. gon. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Neis. gon. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. gon. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Neis. gon. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. gon. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Neis. gon. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. men. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Neis. men. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. men. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Neis. men. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. men. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Neis. men. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. men. Penicillin (measured) | Bacteriology |  |  | Qualit.
S. Neis. men. Penicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Neis. men. Rifampicin (measured) | Bacteriology |  |  | Qualit.
S. Neis. men. Rifampicin (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Penicillin (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Penicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Past. mul. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Past. mul. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Augmentin (Repporting) | Bacteriology |  |  | Qualit.
S. Prot. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Prot. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Cefotaxim (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Prot. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Prot. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Prot. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Prot. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Prot. Co-trimoxazol (Raporting) | Bacteriology |  |  | Qualit.
S. Prot. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Prot. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Prot. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Prot. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Prot. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Prot. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Prov. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Prov. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Prov. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Prov. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Prov. Cefuroxime (Raporting) | Bacteriology |  |  | Qualit.
S. Prov. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Prov. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Prov. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Prov. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Piperacilline/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Prov. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Prov. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Prov. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Ciprofloxacine (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Colistin (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Colistin (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Piperacillin (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Piperacillin (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Pseu. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Pseu. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Salm. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Salm. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Salm. Cefepime (Rapportage) | Bacteriology |  |  | Qualit.
S. Salm. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Salm. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Salm. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Salm. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Cefuroxime (measuerd) | Bacteriology |  |  | Qualit.
S. Salm. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Salm. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Salm. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Salm. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Salm. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Salm. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Salm. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Salm. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Salm. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Ser. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Ser. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Ser. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Ser. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Ser. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Ser. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Ser. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Ser. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Ser. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Ser. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Ser. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Ser. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Ser. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Ser. Tobramycine (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Shig. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Shig. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Shig. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Shig. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Shig. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Shig. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Shig. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Shig. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Shig. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Shig. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Shig. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Shig. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Shig. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Shig. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Shig. Tobramycine (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Clindamycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Clindamycin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Erythromycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Erythromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Flucloxacillin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Flucloxacillin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Rifampicin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Rifampicin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. aur. Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. aur. Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. pseudi. oxacilline (Measured) | Bacteriology |  |  | Qualit.
S. Sta. pseudi. Oxacilline (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Clindamycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Clindamycin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Erythromycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Erythromycin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Flucloxacillin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Flucloxacillin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Rifampicin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Rifampicin (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Tetracyclin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Tetracycline (Reporting) | Bacteriology |  |  | Qualit.
S. Sta. spp. Vancomycin (measured) | Bacteriology |  |  | Qualit.
S. Sta. spp. Vancomycin (Reporting) | Bacteriology |  |  | Qualit.
S. Str. pne. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Str. pne. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Str. pne. Doxycycline (measured) | Bacteriology |  |  | Qualit.
S. Str. pne. Doxycycline (Reporting) | Bacteriology |  |  | Qualit.
S. Str. pne. Penicillin (measured) | Bacteriology |  |  | Qualit.
S. Str. pne. Penicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Vib. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Vib. Cefotaxime (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Vib. Ceftazidime (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Chlooramphenicol (measured) | Bacteriology |  |  | Qualit.
S. Vib. Chlooramphenicol (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Ciprofloxacine (measured) | Bacteriology |  |  | Qualit.
S. Vib. Ciprofloxacine (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Colistine (measured) | Bacteriology |  |  | Qualit.
S. Vib. Colistine (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Vib. Co-trimoxazol (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Erytromycine (measured) | Bacteriology |  |  | Qualit.
S. Vib. Erytromycine (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Vib. Meropenem (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Neomycine (measured) | Bacteriology |  |  | Qualit.
S. Vib. Neomycine (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Piperacilline/ Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Vib. Piperacilline/ Tazobactam (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Tetracycline (measured) | Bacteriology |  |  | Qualit.
S. Vib. Tetracycline (reporting) | Bacteriology |  |  | Qualit.
S. Vib. Trobramycine (measured) | Bacteriology |  |  | Qualit.
S. Vib. Trobramycine (reporting) | Bacteriology |  |  | Qualit.
S. Yer. Amoxicillin (measured) | Bacteriology |  |  | Qualit.
S. Yer. Amoxicillin (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Augmentin (measured) | Bacteriology |  |  | Qualit.
S. Yer. Augmentin (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Cefepime (measured) | Bacteriology |  |  | Qualit.
S. Yer. Cefepime (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Cefotaxime (measured) | Bacteriology |  |  | Qualit.
S. Yer. Cefotaxime (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Ceftazidime (measured) | Bacteriology |  |  | Qualit.
S. Yer. Ceftazidime (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Ceftriaxone (measured) | Bacteriology |  |  | Qualit.
S. Yer. Ceftriaxone (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Cefuroxime (measured) | Bacteriology |  |  | Qualit.
S. Yer. Cefuroxime (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Ciprofloxacin (measured) | Bacteriology |  |  | Qualit.
S. Yer. Ciprofloxacin (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Co-trimoxazol (measured) | Bacteriology |  |  | Qualit.
S. Yer. Co-trimoxazol (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Gentamicin (measured) | Bacteriology |  |  | Qualit.
S. Yer. Gentamicin (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Imipenem (measured) | Bacteriology |  |  | Qualit.
S. Yer. Imipenem (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Meropenem (measured) | Bacteriology |  |  | Qualit.
S. Yer. Meropenem (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Piperacillin/Tazobactam (measured) | Bacteriology |  |  | Qualit.
S. Yer. Piperacillin/Tazobactam (Reporting) | Bacteriology |  |  | Qualit.
S. Yer. Tobramycin (measured) | Bacteriology |  |  | Qualit.
S. Yer. Tobramycin (Reporting) | Bacteriology |  |  | Qualit.
B. henselae IgG | Bartonella henselae serology |  |  | {titer} |  |  |  |  | SDsa1
B. henselae IgM | Bartonella henselae serology |  |  | {titer} |  |  |  |  | SDsa1
B. henselae Igtotaal | Bartonella henselae serology |  |  | {titer} |  |  |  |  | SDsa1
Matches recent Bartonella henselae infection | Bartonella henselae serology |  |  | Qualit.
Clobazam | Benzodiazepines | Based on EMA/FDA | 200 | ug/L | 26.5 | 22.5018 | 53 | 45.0036 | TEa
Clonazepam | Benzodiazepines | Based on EMA/FDA | 60 | ug/L | 26.5 | 29.0077 | 15.9 | 17.4046 | SDsa1
Desmethylclobazam | Benzodiazepines | Based on EMA/FDA | 3000 | ug/L | 26.5 | 15.2343 | 795 | 457.0277 | TEa
Desmethyldiazepam | Benzodiazepines | Based on EMA/FDA | 500 | ug/L | 26.5 | 26.4473 | 132.5 | 132.2367 | TEa
Diazepam | Benzodiazepines | Based on EMA/FDA | 500 | ug/L | 26.5 | 25.2418 | 132.5 | 126.209 | TEa
Hydroxymidazolam | Benzodiazepines | Based on EMA/FDA | 150 | ug/L | 26.5 | 25.1223 | 39.75 | 37.6835 | TEa
Midazolam | Benzodiazepines | Based on EMA/FDA | 200 | ug/L | 26.5 | 25.2154 | 53 | 50.4308 | TEa
Oxazepam | Benzodiazepines | Based on EMA/FDA | 2000 | ug/L | 26.5 | 30.2919 | 530 | 605.8387 | SDsa1
Temazepam | Benzodiazepines | Based on EMA/FDA | 500 | ug/L | 26.5 | 24.4609 | 132.5 | 122.3047 | TEa
Total Bile Acid Salts | Bile acid salts |  | 10 | umol/L |  | 22.0172 |  | 2.2017 | SDsa1
Adalimumab | Biologicals | Based on EMA/FDA | 10 | mg/L | 26.5 | 53.7474 | 2.65 | 5.3747 | SDsa1
Adalimumab antibodies | Biologicals |  |  | Qualit.
Infliximab | Biologicals | Based on EMA/FDA | 5 | mg/L | 26.5 | 41.9591 | 1.325 | 2.098 | SDsa1
Infliximab antibodies | Biologicals |  |  | Qualit.
Vedolizumab | Biologicals | Based on EMA/FDA | 10 | mg/L | 26.5 |  | 2.65 |  | SDsa1
Abn.lymphocyte susp.malignant | Blood Cell Morphology |  |  | Qualit.
Abn.lymphocyte susp.reactive | Blood Cell Morphology |  |  | Qualit.
Agglutination | Blood Cell Morphology |  |  | Qualit.
Alder-Reilly anomaly | Blood Cell Morphology |  |  | Qualit.
Atypical lymphocytes | Blood Cell Morphology |  |  | Qualit.
Auer rods | Blood Cell Morphology |  |  | Qualit.
Basophilic granulocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Blasts | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
cleaved lymfocytes | Blood Cell Morphology |  |  | Qualit.
Döhle bodies | Blood Cell Morphology |  |  | Qualit.
Echinocytes | Blood Cell Morphology |  |  | Qualit.
Elliptocytes | Blood Cell Morphology |  |  | Qualit.
Eosinophilic granulocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Erythroblasts (outside WBC count) | Blood Cell Morphology |  |  | Qualit.
Giant granulocytes | Blood Cell Morphology |  |  | Qualit.
Gumprecht shadows (VHL) | Blood Cell Morphology |  |  | Qualit.
Hairy cells | Blood Cell Morphology |  |  | Qualit.
Howell Jolly bodies | Blood Cell Morphology |  |  | Qualit.
Hypersegmentation | Blood Cell Morphology |  |  | Qualit.
Hypochromasia | Blood Cell Morphology |  |  | Qualit.
Hypogranulation | Blood Cell Morphology |  |  | Qualit.
Hypogranulation | Blood Cell Morphology |  |  | Qualit.
LGL cells > 20% of >2.0x10 | Blood Cell Morphology |  |  | Qualit.
lymfocytes grumelee core pattern | Blood Cell Morphology |  |  | Qualit.
Lymphocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Macrocytosis | Blood Cell Morphology |  |  | Qualit.
Macropolycyte | Blood Cell Morphology |  |  | Qualit.
Macrothrombocytes | Blood Cell Morphology |  |  | Qualit.
Malaria | Blood Cell Morphology |  |  | Qualit.
Maturation dissociation | Blood Cell Morphology |  |  | Qualit.
Megakaryocytes remnants | Blood Cell Morphology |  |  | Qualit.
Megalocytes | Blood Cell Morphology |  |  | Qualit.
Metamyelocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Microcytosis | Blood Cell Morphology |  |  | Qualit.
Monocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Myelocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Neutroph. granulocytes (bands) | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Neutroph. granulocytes (PMNS) | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Not to differentiate | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Plasma cells | Blood Cell Morphology |  |  | %
Plasmablasts | Blood Cell Morphology |  |  | Qualit.
Poikylocytose | Blood Cell Morphology |  |  | Qualit.
Polychromasia | Blood Cell Morphology |  |  | Qualit.
Prolymfocytes | Blood Cell Morphology |  |  | Qualit.
Promonocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Promyelocytes | Blood Cell Morphology |  |  | % |  |  |  |  | Binomial
Pseudo Pelger Huet | Blood Cell Morphology |  |  | Qualit.
Punct.Basophilia | Blood Cell Morphology |  |  | Qualit.
Rouleaux formation | Blood Cell Morphology |  |  | Qualit.
Satellitism | Blood Cell Morphology |  |  | Qualit.
Schizocytes | Blood Cell Morphology |  |  | Qualit.
Sickle cells | Blood Cell Morphology |  |  | Qualit.
SLVL cells | Blood Cell Morphology |  |  | Qualit.
Spherocytes | Blood Cell Morphology |  |  | Qualit.
Stomatocytes | Blood Cell Morphology |  |  | Qualit.
Target cells | Blood Cell Morphology |  |  | Qualit.
Tear drop cells | Blood Cell Morphology |  |  | Qualit.
Thrombocyte aggregates | Blood Cell Morphology |  |  | Qualit.
Toxic granulation | Blood Cell Morphology |  |  | Qualit.
Vacuolisation | Blood Cell Morphology |  |  | Qualit.
P.species antigen rapid test | Blood parasites |  |  | Qualit.
Conclusion malaria antigentest | Blood parasites |  |  | Qualit.
Malaria rapid test | Blood parasites |  |  | Qualit.
Moleculaire sneltest (LAMP) | Blood parasites |  |  | Qualit.
Molecular research Malaria | Blood parasites |  |  | Qualit.
P. falciparum PCR | Blood parasites |  | 20 | {Cq}
P. knowlesi PCR | Blood parasites |  | 20 | {Cq}
P. malariae PCR | Blood parasites |  | 20 | {Cq}
P. ovale PCR | Blood parasites |  | 20 | {Cq}
P. species PCR | Blood parasites |  | 20 | {Cq}
P. vivax PCR | Blood parasites |  | 20 | {Cq}
P.falciparum antigen rapid test | Blood parasites |  |  | Qualit.
Parasite in blood | Blood parasites |  |  | Qualit.
Pf/Pk parasitemie | Blood parasites |  | 5 | % |  | 89.2822 |  | 4.4641 | SDsa1
Bilirubin | Blood pigments in cerebrospinal fluid |  | 0.2 | umol/L | 50 | 122.778 | 0.1 | 0.24556 | SDsa1
Bilirubin (absorption) | Blood pigments in cerebrospinal fluid |  | 100 | {milli_absorbance} | 50 | 172.6816 | 50 | 172.6816 | SDsa1
Most appropriate conclusion with regard to SAB | Blood pigments in cerebrospinal fluid |  |  | Qualit.
Oxyhemoglobin | Blood pigments in cerebrospinal fluid |  | 0.5 | umol/L | 50 | 46.091 | 0.25 | 0.23045 | TEa
Oxyhemoglobin (absorption) | Blood pigments in cerebrospinal fluid |  | 240 | {milli_absorbance} | 50 | 267.8884 | 120 | 642.9321 | SDsa1
Ca++ (actual) | Bloodgas Parameters | Biological variation | 1 | mmol/L | 2.04 | 3.8619 | 0.0204 | 0.038619 | SDsa1
Cl- | Bloodgas Parameters | Biological variation | 100 | mmol/L | 1.47 | 2.492 | 1.47 | 2.492 | SDsa1
Glucose | Bloodgas Parameters | Biological variation | 10 | mmol/L | 6.505 | 7.3781 | 0.6505 | 0.73781 | SDsa1 | 2022-01-01
K+ | Bloodgas Parameters | Biological variation | 5 | mmol/L | 5.607 | 2.7289 | 0.28035 | 0.13644 | TEa
Lactate | Bloodgas Parameters | Biological variation | 2 | mmol/L | 30.939 | 13.2444 | 0.61878 | 0.26489 | TEa
Na+ | Bloodgas Parameters | Biological variation | 140 | mmol/L | 0.725 | 1.7798 | 1.015 | 2.4917 | SDsa1
PCO2 | Bloodgas Parameters | Biological variation | 6.5 | kPa | 5.748 | 8.7409 | 0.37362 | 0.56816 | SDsa1
pH | Bloodgas Parameters | Biological variation | 7.3 |  | 0.223 | 0.36321 | 0.016279 | 0.026514 | SDsa1
PO2 | Bloodgas Parameters | Biological variation | 12.5 | kPa |  | 12.7231 |  | 1.5904 | SDsa1
(clusters of) tumor cells | Bonemarrow morphology |  |  | Qualit.
Aantal megakaryocyten pgz (100x) | Bonemarrow morphology |  |  | Qualit.
Alternative bone marrow diagnosis 1 | Bonemarrow morphology |  |  | Qualit.
Alternative bone marrow diagnosis 1 | Bonemarrow morphology |  |  | Qualit.
Amyloïd | Bonemarrow morphology |  |  | Qualit.
Auer rods | Bonemarrow morphology |  |  | Qualit.
Band form (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Basophil (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Blast incl. promonocyten (bonemarrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Bone marrow particles | Bonemarrow morphology |  |  | Qualit.
Cellularity | Bonemarrow morphology |  |  | Qualit.
Cytogenetics | Bonemarrow morphology |  |  | Qualit.
Cytoplasm. Vacuolisatie | Bonemarrow morphology |  |  | Qualit.
Dysplasie quantitative (Erytropoëse) | Bonemarrow morphology |  |  | Qualit.
Dysplasie quantitative (Granulopoëse) | Bonemarrow morphology |  |  | Qualit.
Dysplasie quantitative (Megakaryopoëse) | Bonemarrow morphology |  |  | Qualit.
Döhle bodies | Bonemarrow morphology |  |  | Qualit.
Eosinophil (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Erytroblast (bonemarrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Haard(en) >50 cellen | Bonemarrow morphology |  |  | Qualit.
Hemofagocytose | Bonemarrow morphology |  |  | Qualit.
Histologiy (biopsy) | Bonemarrow morphology |  |  | Qualit.
Hyperlobulaire vormen | Bonemarrow morphology |  |  | Qualit.
Hypersegmentation | Bonemarrow morphology |  |  | Qualit.
Hypogranulation | Bonemarrow morphology |  |  | Qualit.
Hypolobulaire vormen | Bonemarrow morphology |  |  | Qualit.
Immunophenotyping | Bonemarrow morphology |  |  | Qualit.
Iron incorporation (sideroblasts) | Bonemarrow morphology |  |  | Qualit.
Iron stock | Bonemarrow morphology |  |  | Qualit.
Karyorrhexis | Bonemarrow morphology |  |  | Qualit.
Kernbruggen | Bonemarrow morphology |  |  | Qualit.
Kernmisvormingen | Bonemarrow morphology |  |  | Qualit.
Kwantiteit (Erytropoëse) | Bonemarrow morphology |  |  | Qualit.
Kwantiteit (Granulopoëse) | Bonemarrow morphology |  |  | Qualit.
Kwantiteit (Lymfopoëse) | Bonemarrow morphology |  |  | Qualit.
Kwantiteit (Plasmacellen) | Bonemarrow morphology |  |  | Qualit.
Losliggende kernen | Bonemarrow morphology |  |  | Qualit.
Lymphocyte (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
M-E ratio | Bonemarrow morphology |  |  | {ratio} |  |  |  |  | TE
Macrofagen toegenomen | Bonemarrow morphology |  |  | Qualit.
Meerkernig | Bonemarrow morphology |  |  | Qualit.
Meest waarschijnlijke beenmerg diagnose | Bonemarrow morphology |  |  | Qualit.
Megaloblastair | Bonemarrow morphology |  |  | Qualit.
Metamyelocyte (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Microbiological examination | Bonemarrow morphology |  |  | Qualit.
Micromegakaryocyten | Bonemarrow morphology |  |  | Qualit.
Molecular research | Bonemarrow morphology |  |  | Qualit.
Monocyte (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Monopoiesis (morfology) | Bonemarrow morphology |  |  | Qualit.
Morfologie (Plasmacellen) | Bonemarrow morphology |  |  | Qualit.
Myelocyte (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Neutrophil (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
None | Bonemarrow morphology |  |  | Qualit.
Nuclear malformations | Bonemarrow morphology |  |  | Qualit.
Ongebruikelijke kleine vormen | Bonemarrow morphology |  |  | Qualit.
Ongebruikelijke reuze vormen | Bonemarrow morphology |  |  | Qualit.
Other | Bonemarrow morphology |  |  | Qualit.
Pathologische mestcellen | Bonemarrow morphology |  |  | Qualit.
Percentage promonocyten | Bonemarrow morphology |  |  | % |  |  |  |  | TE
Plasma cell (bone marrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Promyelocyt (bonemarrow) | Bonemarrow morphology |  |  | % |  |  |  |  | Binomial
Pseudo Chédiak Higashi granules | Bonemarrow morphology |  |  | Qualit.
Pseudo Pelger-HUet granules | Bonemarrow morphology |  |  | Qualit.
Pseudo-Gaucher cellen | Bonemarrow morphology |  |  | Qualit.
Representative | Bonemarrow morphology |  |  | Qualit.
Ringsideroblasten % | Bonemarrow morphology |  |  | % |  |  |  |  | TE
Sea blue histiocytes | Bonemarrow morphology |  |  | Qualit.
Suspect malignant | Bonemarrow morphology |  |  | Qualit.
Suspect reactive | Bonemarrow morphology |  |  | Qualit.
Toename mestcellen | Bonemarrow morphology |  |  | Qualit.
Uitrijping (Erytropoëse) | Bonemarrow morphology |  |  | Qualit.
Uitrijping (Granulopoëse) | Bonemarrow morphology |  |  | Qualit.
Bordetella IgG | Bordetella serology |  |  | [IU]/mL |  |  |  |  | TEa
Past bij recente Kinkhoest infectie | Bordetella serology |  |  | Qualit.
Brucella Agglutination | Brucella serology |  |  | IZS ** |  |  |  |  | TEa
Brucella IgG (qual.) | Brucella serology |  |  | Qualit.
Brucella IgG [arb'U]/mL | Brucella serology |  | 1 | [arb'U]/mL |  |  |  |  | SDsa1
Brucella IgG {ratio} | Brucella serology |  | 1 | {ratio} |  |  |  |  | SDsa1
Brucella IgM (qual.) | Brucella serology |  |  | Qualit.
Brucella IgM [arb'U]/mL | Brucella serology |  | 1 | [arb'U]/mL |  |  |  |  | SDsa1
Brucella IgM {ratio} | Brucella serology |  | 1 | {ratio} |  |  |  |  | SDsa1
Other(specify at comments) | Brucella serology |  |  | IZS ** |  |  |  |  | TEa
Past bij infectie met Brucella | Brucella serology |  |  | Qualit.
Busulfan | Busulfan | Based on EMA/FDA | 5 | mg/L | 26.5 | 28.489 | 1.325 | 1.4244 | SDsa1
Amiodarone | Cardiac drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 18.0302 | 0.53 | 0.3606 | TEa
Caffeine | Cardiac drugs | Based on EMA/FDA | 15 | mg/L | 26.5 | 14.4919 | 3.975 | 2.1738 | TEa
Desethylamiodarone | Cardiac drugs | Based on EMA/FDA | 2 | mg/L | 26.5 | 25.857 | 0.53 | 0.51714 | TEa
Digoxin | Cardiac drugs | Based on EMA/FDA | 1.5 | ug/L | 26.5 | 20.0173 | 0.3975 | 0.30026 | TEa
Flecainide | Cardiac drugs | Based on EMA/FDA | 1 | mg/L | 26.5 | 15.6033 | 0.265 | 0.15603 | TEa
Theophylline | Cardiac drugs | Based on EMA/FDA | 10 | mg/L | 26.5 | 14.0345 | 2.65 | 1.4035 | TEa
BNP | Cardiac markers |  | 20 | pmol/L |  | 29.4312 |  | 5.8862 | SDsa1
CKMB act | Cardiac markers | Biological variation | 5 | U/L | 24.073 | 54.4896 | 1.2037 | 2.7245 | SDsa1
CKMB mass | Cardiac markers | Biological variation | 5 | ug/L | 31.157 | 15.3822 | 1.5578 | 0.76911 | TEa
hs-Troponin I | Cardiac markers | Biological variation | 1000 | ng/L | 19.363 | 17.9867 | 193.63 | 179.8668 | TEa | 2023-02-14
hs-Troponin T | Cardiac markers | Clinical decision | 14 | ng/L | 30 | 23.482 | 4.2 | 3.2875 | TEa
Myoglobin | Cardiac markers | Biological variation | 50 | ug/L | 19.643 | 12.4496 | 9.8215 | 6.2248 | TEa
NT-proBNP | Cardiac markers | Biological variation | 200 | pmol/L | 12.967 | 13.0824 | 25.934 | 26.1649 | SDsa1
CDT | CDT | Biological variation | 2 | %{CDT_IFCC} | 15.694 | 32.1605 | 0.31388 | 0.64321 | SDsa1
CDT, qualitative | CDT |  |  | Qualit.
Anti-TG2 IgA ULN categorie | Celiac disease |  |  | Qualit.
IgA against endomysium | Celiac disease |  | 40 | {titer} |  | 171.3069 |  | 68.5227 | SDsa1
IgA against gedeamideerd gliadine | Celiac disease |  | 10 | k[arb'U]/L |  | 24.8837 |  | 2.4884 | SDsa1
IgA against tissue transglutaminase | Celiac disease |  | 25 | k[arb'U]/L |  | 49.9885 |  | 12.4971 | SDsa1
IgG against gedeamideerd gliadine | Celiac disease |  | 10 | k[arb'U]/L |  | 43.7062 |  | 4.3706 | SDsa1
IgG tegen endomysium | Celiac disease |  | 16 | {titer} |  |  |  |  | SDsa1
IgG tegen tissue transglutaminase | Celiac disease |  | 2 | k[arb'U]/L |  | 48.4665 |  | 0.96933 | SDsa1
Donor 1 cellen | Chimerism |  | 25 | % | 30 | 38.5321 | 7.5 | 9.633 | SDsa1
Donor 2 cellen | Chimerism |  | 25 | % | 30 |  | 7.5 |  | TEa
Patiënt cellen | Chimerism |  | 25 | % | 30 | 29.9911 | 7.5 | 7.4978 | TEa
ALAT | Clinical Chemistry, blood | Biological variation | 50 | U/L | 16.08 | 14.4945 | 8.04 | 7.2472 | TEa | 2021-01-01
Albumin | Clinical Chemistry, blood | Biological variation | 40 | g/L | 3.532 | 8.0486 | 1.4128 | 3.2194 | SDsa1 | 2024-01-18
Alk. Phosphatase | Clinical Chemistry, blood | Biological variation | 150 | U/L | 10.541 | 13.1528 | 15.8115 | 19.7293 | SDsa1 | 2022-01-01
Ammonia | Clinical Chemistry, blood |  | 100 | umol/L |  | 18.0195 |  | 18.0195 | SDsa1
Amylase | Clinical Chemistry, blood | Biological variation | 250 | U/L | 13.173 | 9.9744 | 32.9325 | 24.9361 | TEa | 2021-01-01
ASAT | Clinical Chemistry, blood | Biological variation | 50 | U/L | 13.511 | 13.3812 | 6.7555 | 6.6906 | TEa | 2024-01-18
Bilirubin | Clinical Chemistry, blood | Biological variation | 20 | umol/L | 24.565 | 12.9837 | 4.913 | 2.5967 | TEa | 2024-01-18
Bilirubin direct | Clinical Chemistry, blood | Biological variation | 50 | umol/L | 44.547 | 10.9213 | 22.2735 | 5.4607 | TEa
Calcium | Clinical Chemistry, blood | Biological variation | 2.5 | mmol/L | 2.296 | 6.0808 | 0.0574 | 0.15202 | SDsa1 | 2022-01-01
Chloride | Clinical Chemistry, blood | Biological variation | 100 | mmol/L | 1.333 | 4.9584 | 1.333 | 4.9584 | SDsa1 | 2021-01-01
CK | Clinical Chemistry, blood | Biological variation | 150 | U/L | 22.565 | 9.1422 | 33.8475 | 13.7133 | TEa | 2021-01-01
Corrected Calcium | Clinical Chemistry, blood | Biological variation | 2.5 | mmol/L | 2.549 |  | 0.063725 |  | SDsa1
Creatinine | Clinical Chemistry, blood | Biological variation | 100 | umol/L | 7.413 | 8.3015 | 7.413 | 8.3015 | SDsa1 | 2022-01-01
eGFR (F, 55, white) | Clinical Chemistry, blood | Biological variation | 40 | mL/min/{1.73_m2} | 11 | 12.3371 | 4.4 | 4.9348 | SDsa1
Gamma-GT | Clinical Chemistry, blood | Biological variation | 50 | U/L | 18.257 | 9.8993 | 9.4315 | 4.9496 | TEa | 2024-01-18
Glucose | Clinical Chemistry, blood | Biological variation | 10 | mmol/L | 10 Φ | 6.9615 | 1 | 0.69615 | TEa | 2021-01-01
Inorg. Phosphate | Clinical Chemistry, blood | Biological variation | 1 | mmol/L | 9.745 | 8.495 | 0.09745 | 0.08495 | TEa | 2022-01-01
Iron | Clinical Chemistry, blood | Biological variation | 30 | umol/L | 26.654 | 7.194 | 7.9962 | 2.1582 | TEa | 2023-02-14
Lactate | Clinical Chemistry, blood | Biological variation | 1.6 | mmol/L | 30.419 | 8.5333 | 0.4867 | 0.13653 | TEa
LD | Clinical Chemistry, blood | Biological variation | 500 | U/L | 7.698 | 12.8753 | 38.49 | 64.3764 | SDsa1 | 2021-01-01
Lipase | Clinical Chemistry, blood | Biological variation | 25 | U/L | 14.203 | 21.7235 | 3.5507 | 5.4309 | SDsa1 | 2021-01-01
Lithium | Clinical Chemistry, blood | Biological variation | 1 | mmol/L | 12 | 11.3362 | 0.12 | 0.11336 | TEa
Magnesium | Clinical Chemistry, blood | Biological variation | 1 | mmol/L | 4.014 | 7 | 0.04014 | 0.07 | SDsa1 | 2022-01-01
Osmolality | Clinical Chemistry, blood | Biological variation | 300 | mosm/kg | 1.515 | 4.0783 | 4.545 | 12.2348 | SDsa1
Potassium | Clinical Chemistry, blood | Biological variation | 4 | mmol/L | 4.85 | 4.357 | 0.194 | 0.17428 | TEa | 2021-01-01
Pseudo Cholinesterase | Clinical Chemistry, blood | Biological variation | 6 | kU/L | 9.831 | 6.5009 | 0.58986 | 0.39006 | TEa
Sodium | Clinical Chemistry, blood | Biological variation | 140 | mmol/L | 0.692 | 2.876 | 0.9688 | 4.0264 | SDsa1 | 2021-01-01
Total Protein | Clinical Chemistry, blood | Biological variation | 70 | g/L | 3.466 | 5.8259 | 2.4262 | 4.0781 | SDsa1 | 2021-01-01
Urate | Clinical Chemistry, blood | Biological variation | 500 | umol/L | 12.82 | 7.108 | 64.1 | 35.5398 | TEa | 2022-01-01
Urea | Clinical Chemistry, blood | Biological variation | 10 | mmol/L | 17.592 | 8.7925 | 1.7763 | 0.87925 | TEa | 2021-01-01
Acetylacetate | Clinical Chemistry, urine qualitative |  |  | Qualit.
Bilirubin | Clinical Chemistry, urine qualitative |  |  | Qualit.
Erytrocytes/Hb | Clinical Chemistry, urine qualitative |  |  | Qualit.
Glucose | Clinical Chemistry, urine qualitative |  |  | Qualit.
Ketones | Clinical Chemistry, urine qualitative |  |  | Qualit.
Leucocytes | Clinical Chemistry, urine qualitative |  |  | Qualit.
Nitrite | Clinical Chemistry, urine qualitative |  |  | Qualit.
pH | Clinical Chemistry, urine qualitative |  |  | Qualit.
Pregnancy | Clinical Chemistry, urine qualitative |  |  | Qualit.
Protein | Clinical Chemistry, urine qualitative |  |  | Qualit.
S.G. | Clinical Chemistry, urine qualitative |  |  | Qualit.
Urobilinogen | Clinical Chemistry, urine qualitative |  |  | Qualit.
(micro)Albumin | Clinical Chemistry, urine quantitative | Biological variation | 100 | mg/L | 46.134 | 14.6706 | 46.134 | 14.6706 | TEa
Amylase | Clinical Chemistry, urine quantitative | Biological variation | 200 | U/L | 103.713 | 17.6399 | 207.426 | 35.2797 | TEa
Calcium | Clinical Chemistry, urine quantitative | Biological variation | 1 | mmol/L | 34.133 | 13.2772 | 0.34133 | 0.13277 | TEa
Chloride | Clinical Chemistry, urine quantitative |  | 100 | mmol/L |  | 8.3928 |  | 8.3928 | SDsa1
Creatinine | Clinical Chemistry, urine quantitative | Biological variation | 6 | mmol/L | 28.374 | 9.6052 | 1.7024 | 0.57631 | TEa
Glucose | Clinical Chemistry, urine quantitative |  | 10 | mmol/L |  | 9.6498 |  | 0.96498 | SDsa1
Inorg. Phosphate | Clinical Chemistry, urine quantitative | Biological variation | 15 | mmol/L | 22.073 | 10.7796 | 3.311 | 1.6169 | TEa
Magnesium | Clinical Chemistry, urine quantitative | Biological variation | 4 | mmol/L | 45.015 | 12.6836 | 1.8006 | 0.50734 | TEa
Osmolality | Clinical Chemistry, urine quantitative | Biological variation | 450 | mosm/kg | 39.459 | 2.8994 | 177.5655 | 13.0474 | TEa
Oxalate | Clinical Chemistry, urine quantitative | Biological variation | 0.1 | mmol/L | 46.795 |  | 0.046795 |  | SDsa1
Potassium | Clinical Chemistry, urine quantitative | Biological variation | 60 | mmol/L | 28.377 | 8.1234 | 17.0262 | 4.8741 | TEa
Sodium | Clinical Chemistry, urine quantitative | Biological variation | 100 | mmol/L | 31.979 | 5.5197 | 31.979 | 5.5197 | TEa
Total Protein | Clinical Chemistry, urine quantitative | Biological variation | 1000 | mg/L | 39.959 | 14.9059 | 399.59 | 149.0591 | TEa
Urate | Clinical Chemistry, urine quantitative | Biological variation | 1000 | mmol/L | 19.392 | 13.1688 | 0.19392 | 0.13169 | TEa
Urea | Clinical Chemistry, urine quantitative | Biological variation | 200 | mmol/L | 22.052 | 10.8632 | 44.104 | 21.7265 | TEa
CMV IgG | CMV/EBV/Toxoplasma serology |  |  | [arb'U]/mL |  |  |  |  | TEa
CMV IgM | CMV/EBV/Toxoplasma serology |  |  | {ratio} |  |  |  |  | TEa
CMV-IgG aviditeit | CMV/EBV/Toxoplasma serology |  |  | {ratio} |  |  |  |  | TEa
EBNA IgG | CMV/EBV/Toxoplasma serology |  |  | [arb'U]/mL |  |  |  |  | TEa
EBV-EA IgG | CMV/EBV/Toxoplasma serology |  |  | [arb'U]/mL |  |  |  |  | TEa
EBV-EA IgM | CMV/EBV/Toxoplasma serology |  |  | {OD_unit} |  |  |  |  | TEa
EBV-VCA IgG | CMV/EBV/Toxoplasma serology |  |  | [arb'U]/mL |  |  |  |  | TEa
EBV-VCA IgM | CMV/EBV/Toxoplasma serology |  |  | [arb'U]/mL |  |  |  |  | TEa
Heterof a.s./slide tst/Paul&Bunell | CMV/EBV/Toxoplasma serology |  |  | N/A |  |  |  |  | TEa
Other(specify at comments) | CMV/EBV/Toxoplasma serology |  |  | {other} |  |  |  |  | TEa
Past bij doorgemaakte CMV-inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Past bij doorgemaakte EBV-inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Past bij doorgemaakte Toxoplasma inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Past bij recente CMV-inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Past bij recente EBV-inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Past bij recente Toxoplasma-inf. | CMV/EBV/Toxoplasma serology |  |  | Qualit.
Toxoplasma IgG | CMV/EBV/Toxoplasma serology |  |  | [IU]/mL |  |  |  |  | TEa
Toxoplasma IgG aviditeit | CMV/EBV/Toxoplasma serology |  |  | {ratio} |  |  |  |  | TEa
Toxoplasma IgM | CMV/EBV/Toxoplasma serology |  |  | {ratio} |  |  |  |  | TEa
Antitrombin | Coagulation | Biological variation | 100 | %{activity} | 4.932 | 13.0818 | 4.932 | 13.0818 | SDsa1 | 2024-01-30
APTT | Coagulation | Biological variation | 50 | s | 4.481 | 9.6731 | 2.2405 | 4.8366 | SDsa1 | 2024-01-30
Factor VIII | Coagulation | Biological variation | 100 | %{activity} | 13.208 | 13.208 | 8.883 | 21.5792 | SDsa1 | 2024-01-30
Fibrinogen | Coagulation | Biological variation | 3 | g/L | 13.598 | 17.0179 | 0.40794 | 0.51054 | SDsa1
PT | Coagulation | Biological variation | 20 | s | 3.576 | 12.7823 | 0.7152 | 2.5565 | SDsa1 | 2024-01-30
PT-INR | Coagulation | Biological variation | 1 | {INR} | 3.371 | 12.6375 | 0.03371 | 0.12637 | SDsa1 | 2024-01-30
ANA (HEp2 IIF screen) | Collagen |  | 500 | {titer} |  | 459.9058 |  | 2299.5291 | SDsa1
antibodies against dsDNA | Collagen |  | 100 | [IU]/mL |  | 74.5542 |  | 74.5542 | SDsa1
Antibodies against ENA (solid phase screen) | Collagen |  |  | Qualit.
CenpB | Collagen |  |  | Qualit.
Dominant HEp2 IIF pattern | Collagen |  |  | Qualit.
Histonen | Collagen |  |  | Qualit.
Jo-1 | Collagen |  |  | Qualit.
Number of HEp2 IIF patterns | Collagen |  |  | Qualit.
Ribosomaal P protein (rRNP) | Collagen |  |  | Qualit.
RNP | Collagen |  |  | Qualit.
Sm | Collagen |  |  | Qualit.
SS-A (52 + 60 kD) | Collagen |  |  | Qualit.
SS-A (52kD) | Collagen |  |  | Qualit.
SS-A (60 kD) | Collagen |  |  | Qualit.
SS-B | Collagen |  |  | Qualit.
Topoisomerase I (Scl70) | Collagen |  |  | Qualit.
AP | Complement factors | Biological variation | 60 | % | 40.25 | 40.5793 | 24.15 | 24.3476 | SDsa1
C1q | Complement factors | Biological variation | 200 | mg/L | 47.25 | 44.4656 | 94.5 | 88.9312 | TEa
C3 | Complement factors | Biological variation | 1.8 | g/L | 7.765 | 12.2721 | 0.13977 | 0.2209 | SDsa1 | 2021-01-01
C4 | Complement factors | Biological variation | 0.35 | g/L | 12.056 | 12.2576 | 0.042196 | 0.042902 | SDsa1 | 2021-01-01
CP | Complement factors | Biological variation | 60 | % | 34.9 | 44.434 | 20.94 | 26.6604 | SDsa1
MBL | Complement factors |  | 1.5 | mg/L |  |  |  |  | SDsa1
MBL-P | Complement factors |  | 50 | % |  | 63.5149 |  | 31.7575 | SDsa1
antibodies against GAD | Diabetes |  | 50 | k[arb'U]/L |  | 234.2607 |  | 117.1303 | SDsa1
antibodies against IA2 | Diabetes |  | 200 | k[arb'U]/L |  | 206.9777 |  | 413.9554 | SDsa1
antibodies against ICA | Diabetes |  | 50 | {titer} |  |  |  |  | SDsa1
Acenocoumarol | Drugs alternating | Based on EMA/FDA | 0.1 | mg/L | 26.5 |  | 0.0265
Adalimumab | Drugs alternating | Based on EMA/FDA | 10 | mg/L | 26.5 |  | 2.65
Aloe-emodin | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Amoxicillin | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
Atenolol | Drugs alternating | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5
Baclofen | Drugs alternating | Based on EMA/FDA | 512 | ug/L | 26.5 |  | 135.68
Benzylpenicillin | Drugs alternating | Based on EMA/FDA | 20 | mg/L | 26.5 |  | 5.3
Bisacodyldiphenol | Drugs alternating | Based on EMA/FDA | 3 | mg/L | 26.5 |  | 0.795
Bisoxantindiphenol | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Bupropion | Drugs alternating | Based on EMA/FDA | 40 | ug/L | 26.5 |  | 10.6
Cefuroxime | Drugs alternating | Based on EMA/FDA | 2 | mg/L | 26.5 |  | 0.53
Chloortalidon | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Chlorothiazide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Chrysophanic acid | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Clopentixol | Drugs alternating | Based on EMA/FDA | 10 | ug/L | 26.5 | 28.5208 | 2.65 | 2.8521
Danthron | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Desmethylflunitrazepam | Drugs alternating | Based on EMA/FDA | 12 | ug/L | 26.5 |  | 3.18
Desmethylmaprotiline | Drugs alternating | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5
Disopyramide | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
Duloxetine | Drugs alternating | Based on EMA/FDA | 80 | ug/L | 26.5 |  | 21.2
Emodin | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Felbamaat | Drugs alternating | Based on EMA/FDA | 55 | mg/L | 26.5 |  | 14.575
Fenprocoumon | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Flucloxacillin | Drugs alternating | Based on EMA/FDA | 100 | mg/L | 26.5 |  | 26.5
Flunitrazepam | Drugs alternating | Based on EMA/FDA | 12 | ug/L | 26.5 | 27.973 | 3.18 | 3.3568
Flupenthixol | Drugs alternating | Based on EMA/FDA | 6 | ug/L | 26.5 |  | 1.59
Furosemide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Gabapentine | Drugs alternating | Based on EMA/FDA | 7.8 | mg/L | 26.5 |  | 2.067
Glibenclamide | Drugs alternating | Based on EMA/FDA | 0.15 | mg/L | 26.5 |  | 0.03975
Haloperidol | Drugs alternating | Based on EMA/FDA | 10 | ug/L | 26.5 | 23.736 | 2.65 | 2.3736
Hydrochlororthiazide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Hydroxybupropion | Drugs alternating | Based on EMA/FDA | 700 | ug/L | 26.5 |  | 185.5
Hydroxychloroquine | Drugs alternating | Based on EMA/FDA | 100 | mg/L | 26.5 |  | 26.5
Hydroxymidazolam | Drugs alternating | Based on EMA/FDA | 150 | ug/L | 26.5 | 25.6092 | 39.75 | 38.4138
Hydroxymidazolamglucuronide | Drugs alternating | Based on EMA/FDA | 500 | ug/L | 26.5 |  | 132.5
Infliximab | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
Kinidine | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
Lacosamide | Drugs alternating | Based on EMA/FDA | 5.3 | mg/L | 26.5 |  | 1.4045
Maprotiline | Drugs alternating | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5
mefruside | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
meropenem | Drugs alternating | Based on EMA/FDA | 5.5 | mg/L | 26.5 |  | 1.4575
Metformin | Drugs alternating | Based on EMA/FDA | 2 | mg/L | 26.5 |  | 0.53
Metoprolol | Drugs alternating | Based on EMA/FDA | 350 | ug/L | 26.5 | 50.4607 | 92.75 | 176.6126
Midazolam | Drugs alternating | Based on EMA/FDA | 200 | ug/L | 26.5 | 24.404 | 53 | 48.808
moclobemide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Mycophenolic acid | Drugs alternating | Based on EMA/FDA | 10.5 | mg/L | 26.5 |  | 2.7825
N4-acetyl-sulfamethoxazole | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
N-acetyl-procainamide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Penfluridol | Drugs alternating | Based on EMA/FDA | 7 | ug/L | 26.5 |  | 1.855
Phenolphtaleïne | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Pimozide | Drugs alternating | Based on EMA/FDA | 12.5 | ug/L | 26.5 |  | 3.3125
Pipamperon | Drugs alternating | Based on EMA/FDA | 400 | ug/L | 26.5 |  | 106
Pregabaline | Drugs alternating | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325
Procaïnamide | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Propranolol | Drugs alternating | Based on EMA/FDA | 150 | ug/L | 26.5 | 6.1685 | 39.75 | 9.2527
Rhein | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Sertindol | Drugs alternating | Based on EMA/FDA | 150 | ug/L | 26.5 |  | 39.75
Sotalol | Drugs alternating | Based on EMA/FDA | 1900 | ug/L | 26.5 |  | 503.5
Spironolacton | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Sulfamethoxazole | Drugs alternating | Based on EMA/FDA | 60 | mg/L | 26.5 |  | 15.9
Teicoplanine | Drugs alternating | Based on EMA/FDA |  | mg/L | 26.5
Tolbutamide | Drugs alternating | Based on EMA/FDA | 50 | mg/L | 26.5 |  | 13.25
Topiramate | Drugs alternating | Based on EMA/FDA | 12 | mg/L | 26.5 |  | 3.18
tranylcypromine | Drugs alternating | Based on EMA/FDA | 30 | ug/L | 26.5 |  | 7.95
Triamterene | Drugs alternating | Based on EMA/FDA | 1 | mg/L | 26.5 |  | 0.265
Trimethoprim | Drugs alternating | Based on EMA/FDA | 2 | mg/L | 26.5 |  | 0.53
Warfarin | Drugs alternating | Based on EMA/FDA | 2 | mg/L | 26.5 |  | 0.53
Acetaminophen confirmation | Drugs of Abuse | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325 |  | SDA
Acetaminophen screening | Drugs of Abuse | Based on EMA/FDA | 5 | mg/L | 26.5 |  | 1.325 |  | SDA
Alfa-PVP screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Alfa-PVP confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Alprazolam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Amphetamine confirmation | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Amphetamine screening | Drugs of Abuse | Based on EMA/FDA | 500 | ug/L | 26.5 |  | 132.5 |  | SDA
Aspect urine | Drugs of Abuse | Based on EMA/FDA |  | Qualit.
Barbiturate confirmation | Drugs of Abuse | Based on EMA/FDA | 150 | ug/L | 26.5 |  | 39.75 |  | SDA
Barbiturate screening | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Benzodiazepine screening | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Benzoylecgonine confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Bromazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Buprenorphine confirmation | Drugs of Abuse | Based on EMA/FDA | 2 | ug/L | 26.5 |  | 0.53 |  | SDA
Buprenorphine screening | Drugs of Abuse | Based on EMA/FDA | 5 | ug/L | 26.5 |  | 1.325 |  | SDA
Cannabis screeninng | Drugs of Abuse | Based on EMA/FDA | 50 | ug/L | 26.5 |  | 13.25 |  | SDA
Clonazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Cocaïne confirmation | Drugs of Abuse | Based on EMA/FDA | 150 | ug/L | 26.5 |  | 39.75 |  | SDA
Cocaïne screening | Drugs of Abuse | Based on EMA/FDA | 150 | ug/L | 26.5 |  | 39.75 |  | SDA
Codeine confirmation | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
Creatinine | Drugs of Abuse | Biological variation | 2 | mmol/L | 28.374 |  | 0.56748 |  | SDA
Diazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Drugs | Drugs of Abuse | Based on EMA/FDA |  | Qualit.
EDDP confirmation | Drugs of Abuse | Based on EMA/FDA | 75 | ug/L | 26.5 |  | 19.875 |  | SDA
EDDP screening | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Ethanol confirmation | Drugs of Abuse | Based on EMA/FDA | 0.1 | g/L | 26.5 |  | 0.0265 |  | SDA
Ethanol screening | Drugs of Abuse | Based on EMA/FDA | 0.1 | g/L | 26.5 |  | 0.0265 |  | SDA
Ethylglucuroninde confirmation | Drugs of Abuse | Based on EMA/FDA | 500 | ug/L | 26.5 |  | 132.5 |  | SDA
Ethylglucuroninde screening | Drugs of Abuse | Based on EMA/FDA | 500 | ug/L | 26.5 |  | 132.5 |  | SDA
Fentanyl screening | Drugs of Abuse | Based on EMA/FDA | 1 | ug/L | 26.5 |  | 0.265 |  | SDA
Fentanyl confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Flunitrazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Flurazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
GHB screening | Drugs of Abuse | Based on EMA/FDA | 10 | mg/L | 26.5 |  | 2.65 |  | SDA
GHB confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | mg/L | 26.5 |  | 2.65 |  | SDA
Ketamine confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Ketamine screening | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Lorazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Lormetazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
LSD confirmation | Drugs of Abuse | Based on EMA/FDA | 1 | ug/L | 26.5 |  | 0.265 |  | SDA
LSD screening | Drugs of Abuse | Based on EMA/FDA | 1 | ug/L | 26.5 |  | 0.265 |  | SDA
MDA confirmation | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
MDEA confirmation | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
MDMA confirmation | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Methadone confirmation | Drugs of Abuse | Based on EMA/FDA | 250 | ug/L | 26.5 |  | 66.25 |  | SDA
Methadone screening | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
Methylamphetamine confirmation | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Methylfenidaat screening | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Methylfenidaat Confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Midazolam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Mitragynine screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Mitragynine confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Morphine confirmation | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
Nitrazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Norcodeïne confirmation | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
Nordiazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Opiates screening | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
Oxazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Phenazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
Phencyclidine screening | Drugs of Abuse | Based on EMA/FDA | 25 | ug/L | 26.5 |  | 6.625 |  | SDA
Phencyclidine confirmation | Drugs of Abuse | Based on EMA/FDA | 25 | ug/L | 26.5 |  | 6.625 |  | SDA
Propoxyfeen screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
Propoxyfeen confirmation | Drugs of Abuse | Based on EMA/FDA | 300 | ug/L | 26.5 |  | 79.5 |  | SDA
TCA confirmation | Drugs of Abuse | Based on EMA/FDA | 1000 | ug/L | 26.5 |  | 265 |  | SDA
TCA screening | Drugs of Abuse | Based on EMA/FDA | 1000 | ug/L | 26.5 |  | 265 |  | SDA
Temazepam confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
THCCOOH confirmation | Drugs of Abuse | Based on EMA/FDA | 15 | ug/L | 26.5 |  | 3.975 |  | SDA
Tramadol screening | Drugs of Abuse | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDA
Tramadol confirmation | Drugs of Abuse | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDA
XTC/methamfetaminen screening | Drugs of Abuse | Based on EMA/FDA | 500 | ug/L | 26.5 |  | 132.5 |  | SDA
3-MMC screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
3-MMC confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
4-FA screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
4-FA confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
4-MMC/mefedron screening | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
4-MMC/mefedron confirmation | Drugs of Abuse | Based on EMA/FDA | 10 | ug/L | 26.5 |  | 2.65 |  | SDA
3 celspanel | Erythrocyte serology basic |  |  | Qualit.
ABO conclusie | Erythrocyte serology basic |  |  | Qualit.
Anti-A | Erythrocyte serology basic |  |  | Qualit.
Anti-B | Erythrocyte serology basic |  |  | Qualit.
Anti-D | Erythrocyte serology basic |  |  | Qualit.
Anti-D controle | Erythrocyte serology basic |  |  | Qualit.
A-testerytrocytes | Erythrocyte serology basic |  |  | Qualit.
B-testerytrocytes | Erythrocyte serology basic |  |  | Qualit.
Kruis. donor 1 | Erythrocyte serology basic |  |  | Qualit.
Kruis. donor 2 | Erythrocyte serology basic |  |  | Qualit.
Rhesus-D conclusie | Erythrocyte serology basic |  |  | Qualit.
Anti C | Erythrocyte serology detailed |  |  | Qualit.
Anti c | Erythrocyte serology detailed |  |  | Qualit.
Anti C(w) | Erythrocyte serology detailed |  |  | Qualit.
Anti D | Erythrocyte serology detailed |  |  | Qualit.
Anti E | Erythrocyte serology detailed |  |  | Qualit.
Anti e | Erythrocyte serology detailed |  |  | Qualit.
Anti f | Erythrocyte serology detailed |  |  | Qualit.
Anti Fy(a) | Erythrocyte serology detailed |  |  | Qualit.
Anti Fy(b) | Erythrocyte serology detailed |  |  | Qualit.
Anti Jk(a) | Erythrocyte serology detailed |  |  | Qualit.
Anti Jk(b) | Erythrocyte serology detailed |  |  | Qualit.
Anti K | Erythrocyte serology detailed |  |  | Qualit.
Anti k | Erythrocyte serology detailed |  |  | Qualit.
Anti Kp(a) | Erythrocyte serology detailed |  |  | Qualit.
Anti Le(a) | Erythrocyte serology detailed |  |  | Qualit.
Anti Le(b) | Erythrocyte serology detailed |  |  | Qualit.
Anti Lu(a) | Erythrocyte serology detailed |  |  | Qualit.
Anti M | Erythrocyte serology detailed |  |  | Qualit.
Anti N | Erythrocyte serology detailed |  |  | Qualit.
Anti P1 | Erythrocyte serology detailed |  |  | Qualit.
Anti S | Erythrocyte serology detailed |  |  | Qualit.
Anti s | Erythrocyte serology detailed |  |  | Qualit.
Anti Wr(a) | Erythrocyte serology detailed |  |  | Qualit.
3 celspanel | Erythrocyte serology pregnancy |  |  | Qualit.
ABO conclusie | Erythrocyte serology pregnancy |  |  | Qualit.
Anti-A | Erythrocyte serology pregnancy |  |  | Qualit.
Anti-B | Erythrocyte serology pregnancy |  |  | Qualit.
Anti-c | Erythrocyte serology pregnancy |  |  | Qualit.
Anti-D | Erythrocyte serology pregnancy |  |  | Qualit.
Anti-D controle | Erythrocyte serology pregnancy |  |  | Qualit.
A-testerytrocytes | Erythrocyte serology pregnancy |  |  | Qualit.
B-testerytrocytes | Erythrocyte serology pregnancy |  |  | Qualit.
Rhesus-c conclusion | Erythrocyte serology pregnancy |  |  | Qualit.
Rhesus-D conclusie | Erythrocyte serology pregnancy |  |  | Qualit.
Ethanol | Ethanol in serum |  | 2.5 | [ppth] |  | 8.9534 |  | 0.22384 | SDsa1
Doorgemaakte/vaccinatie Mazelen inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Doorgemaakte/vaccinatie Rubella inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Herpes simplex antibodies | Exantheme Diseases Diagnostics |  |  | {titer}
Herpes simplex IgG | Exantheme Diseases Diagnostics |  |  | IZS **
Herpes simplex IgM | Exantheme Diseases Diagnostics |  |  | IZS **
Herpes simplex type 2 IgG | Exantheme Diseases Diagnostics |  |  | IZS **
Herpes simplex type 2 IgM | Exantheme Diseases Diagnostics |  |  | IZS **
HHV-6 IgG | Exantheme Diseases Diagnostics |  |  | {titer}
HHV-6 IgM | Exantheme Diseases Diagnostics |  |  | {titer}
Measels antibodies | Exantheme Diseases Diagnostics |  |  | {titer}
Measels IgG | Exantheme Diseases Diagnostics |  |  | IZS **
Measels IgM | Exantheme Diseases Diagnostics |  |  | IZS **
Other(specify at comments) | Exantheme Diseases Diagnostics |  |  | IZS **
Parvovirus B19 IgG | Exantheme Diseases Diagnostics |  |  | IZS **
Parvovirus B19 IgM | Exantheme Diseases Diagnostics |  |  | IZS **
Past bij doorgemaakte HSV inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij doorgemaakte parvo B19 inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij doorgemaakte VZV inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij recente HSV inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij recente Mazelen inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij recente Parvo. B19 inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij recente Rubella inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Past bij recente VZV inf. | Exantheme Diseases Diagnostics |  |  | Qualit.
Rubella IgG | Exantheme Diseases Diagnostics |  |  | [IU]/mL
Rubella IgM | Exantheme Diseases Diagnostics |  |  | IZS **
Varicella zoster antibodies | Exantheme Diseases Diagnostics |  |  | {titer}
Varicella zoster IgG | Exantheme Diseases Diagnostics |  |  | IZS **
Varicella zoster IgM | Exantheme Diseases Diagnostics |  |  | IZS **
Alpha-1-antitrypsin | Feces-chemistry |  | 17 | mg/g |  | 100.1554 |  | 17.0264 | SDsa1
Bile acids | Feces-chemistry |  | 5 | umol/g |  |  |  |  | SDsa1
Calprotectin | Feces-chemistry |  | 60 | ug/g |  | 58.3344 |  | 35.0007 | SDsa1
Elastase | Feces-chemistry |  | 275 | ug/g |  | 55.8246 |  | 153.5176 | SDsa1
Fat in feces | Feces-chemistry |  | 10 | g/(100.g) |  | 50.1735 |  | 5.0174 | SDsa1
Fetal fibronectin | Fetal fibronectine |  | 250 | ng/mL |  | 45.3015 |  | 113.2538 | SDsa1
Fetal cells | Fetomaternal transfusion |  | 1 | % |  | 50.0413 |  | 0.50041 | SDsa1
Foetomaternale hemorrhage | Fetomaternal transfusion |  | 50 | mL |  | 46.5953 |  | 23.2977 | SDsa1
Fructosamin | Fructosamin | Biological variation | 300 | umol/L | 3.574 | 8.6846 | 10.722 | 26.0539 | SDsa1 | 2021-01-01
HbA | Hb variants | Biological variation | 50 | % | 29.28 | 25.0737 | 14.64 | 12.5369 | TEa
HbA2 | Hb variants | Biological variation | 3 | % | 50 | 37.1459 | 1.5 | 1.1144 | Tea
HbF | Hb variants | Biological variation | 3 | % | 45 | 34.8489 | 1.35 | 1.0455 | TEa
HbX1 | Hb variants | Biological variation | 50 | % | 18.3 | 17.6819 | 9.15 | 8.841 | TEa
HbX2 | Hb variants | Biological variation | 0.1 | % | 330 | 15.7006 | 0.015701 | 0.17127 | TEa
HbA1c | HbA1c whole blood | Biological variation | 50 | mmol/mol | 3.14 | 8.0643 | 1.57 | 4.0322 | SDsa1 | 2022-01-01
Helicobacter pylori Antigeen in Feces | Helicobacter pylori (Ag) in feces |  |  | N/A
Helicobacter antibodies | Helicobacter serology |  |  | {titer}
Helicobacter IgA | Helicobacter serology |  |  | IZS **
Helicobacter IgG | Helicobacter serology |  |  | IZS **
Helicobacter serologie past bij | Helicobacter serology |  |  | Qualit.
Other(specify at comments) | Helicobacter serology |  |  | IZS **
Ancylostoma spp PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Ascaris spp PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Necator americanus PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Schistosoma spp PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Strongyloides stercoralis PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Trichuris trichiura PCR | Helminth External Mol. Quality Assessment |  | 20 | {Cq} |  |  |  |  | SDsa1
Erytrocytes | Hemocytometry | Biological variation | 4 | 10*12/L | 4.195 | 3.8717 | 0.1678 | 0.15487 | TEa | 2024-01-16
Hematocrit | Hemocytometry | Biological variation | 0.4 | L/L | 3.875 | 5.8681 | 0.0155 | 0.023473 | SDsa1 | 2024-01-16
Hemoglobin | Hemocytometry | Biological variation | 7 | mmol/L | 3.895 | 3.7917 | 0.27265 | 0.26542 | TEa | 2024-01-16
IPF | Hemocytometry | Biological variation | 3.2 | % | 10 | 56.8431 | 0.32 | 1.819 | SDsa1
Leucocytes | Hemocytometry | Biological variation | 5 | 10*9/L | 14.191 | 7.8478 | 0.70955 | 0.39239 | TEa | 2024-01-16
Lymfocytes abs. | Hemocytometry | Biological variation | 1.5 | 10*9/L | 15.195 | 12.4054 | 0.22792 | 0.18608 | TEa | 2021-01-01
MCV | Hemocytometry | Biological variation | 100 | fL | 1.631 | 4.2804 | 1.631 | 4.2804 | SDsa1 | 2024-01-16
Monocytes abs. | Hemocytometry | Biological variation | 0.4 | 10*9/L | 17.442 | 23.6142 | 0.069768 | 0.094457 | SDsa1 | 2021-01-01
Neutrofiles abs. | Hemocytometry | Biological variation | 3 | 10*9/L | 18.41 | 8.8254 | 0.5523 | 0.26476 | TEa | 2021-01-01
RDW-SD | Hemocytometry | Biological variation | 35 | fL | 2.558 | 9.4896 | 0.8953 | 3.3214 | SDsa1 | 2021-01-01
RET-He | Hemocytometry | Biological variation | 2700 | amol | 2.447 | 6.2822 | 66.069 | 169.6202 | SDsa1 | 2023-01-01
Reticulocytes abs. | Hemocytometry | Biological variation | 50 | 10*9/L | 15.198 | 28.9647 | 7.599 | 14.4823 | SDsa1 | 2021-01-01
Trombocytes | Hemocytometry | Biological variation | 175 | 10*9/L | 10.216 | 12.1598 | 17.878 | 21.2797 | SDsa1 | 2024-01-16
Fraction COHb | Hemoximetry |  | 15 | % |  | 9.0821 |  | 1.3623 | SDsa1
Fraction HHb | Hemoximetry |  | 1 | % |  |  |  |  | SDsa1
Fraction MetHb | Hemoximetry |  | 20 | % |  | 7.6832 |  | 1.5366 | SDsa1
Fraction O2Hb | Hemoximetry |  | 90 | % |  | 3.0229 |  | 2.7206 | SDsa1
sO2 | Hemoximetry |  | 50 | % |  | 4.3167 |  | 2.1583 | SDsa1
Total Hb | Hemoximetry | Biological variation | 8 | mmol/L | 3.85 | 6.9418 | 0.308 | 0.55534 | SDsa1 | 2021-01-01
Beschermd tegen hepatitis A? | Hepatitis ABCE serology |  |  | Qualit.
Beschermd tegen hepatitis B? | Hepatitis ABCE serology |  |  | Qualit.
Hepatitis A IgG | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis A IgM | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis A IgT | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B Anti-HBcore | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B Anti-HBe | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B Anti-HBs | Hepatitis ABCE serology |  |  | [IU]/L
Hepatitis B HBeAg | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B HBsAg | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B HBsAg confirmation | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis B HBsAg kwantitatief | Hepatitis ABCE serology |  |  | [IU]/mL
Hepatitis B IgM anti-HBcore | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis C anti C1 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti c100p + 5-1-1p | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti C2 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti c22p | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti C33c | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti NS3 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti NS4 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C anti NS5 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis C Anti-HCV | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis C Anti-HCV confirmation | Hepatitis ABCE serology |  |  | Qualit.
Hepatitis C-anti E2 | Hepatitis ABCE serology |  |  | IZS **
Hepatitis E IgG | Hepatitis ABCE serology |  |  | {ratio}
Hepatitis E IgM | Hepatitis ABCE serology |  |  | {ratio}
Other(specify at comments) | Hepatitis ABCE serology |  |  | IZS **
past bij acute infectie hep A | Hepatitis ABCE serology |  |  | Qualit.
Past bij acute infectie hep E | Hepatitis ABCE serology |  |  | Qualit.
past bij acute/chronisch inf. hep B | Hepatitis ABCE serology |  |  | Qualit.
past bij doorgemaakte hep B inf. | Hepatitis ABCE serology |  |  | Qualit.
Past bij infectie hep C | Hepatitis ABCE serology |  |  | Qualit.
anti HIV-gp105 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-gp120 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-gp160 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-gp36 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-gp41 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p17 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p24 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p31 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p39 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p51 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p55 | HIV serology (incl. blot) |  |  | IZS **
anti HIV-p66 | HIV serology (incl. blot) |  |  | IZS **
Fits infection with HIV | HIV serology (incl. blot) |  |  | Qualit.
HIV p24 antigentest | HIV serology (incl. blot) |  |  | Qualit.
HIV screening (1e test) | HIV serology (incl. blot) |  |  | {ratio}
HIV screening (2e test) | HIV serology (incl. blot) |  |  | {ratio}
HIV-confirmation (immuneblot) | HIV serology (incl. blot) |  |  | IZS **
Past bij acute infectie met HIV | HIV serology (incl. blot) |  |  | Qualit.
Homocystein | Homocystein | Biological variation | 1 | nmol/L | 15.476 | 21.6311 | 3.0952 | 4.3262 | SDsa1
11-Deoxycortisole | Hormones in Saliva | Biological variation | 1 | nmol/L | 27.68 |  | 0.2768 |  | SDsa1
21-deoxycortisole | Hormones in Saliva |  | 1 | nmol/L |  |  |  |  | SDsa1
Androsteendion | Hormones in Saliva | Biological variation | 5 | nmol/L | 22.672 |  | 0.22672 |  | SDsa1
Cortisol | Hormones in Saliva | Biological variation | 20 | nmol/L |  | 37.2134 |  | 1.8607 | SDsa1
Cortison | Hormones in Saliva | Biological variation | 2 | nmol/L |  | 24.0349 |  | 4.807 | SDsa1
DHEA | Hormones in Saliva | Biological variation | 1 | ng/L | 27.68 |  | 0.5536 |  | SDsa1
Melatonin | Hormones in Saliva | Biological variation | 3 | nmol/L |  |  |  |  | SDsa1
Progesteron,17OH | Hormones in Saliva | Biological variation | 1 | nmol/L | 27.68 |  | 0.8304 |  | SDsa1
Testosterone | Hormones in Saliva | Biological variation |  |  | 22.672 |  | 0.22672 |  | SDsa1
11 deoxycorticosterone | Hormones in Serum |  | 0.5 | nmol/L |  |  |  |  | TEa
11-Deoxycortisol | Hormones in Serum | Biological variation | 200 | nmol/L | 27.079 | 27.7417 | 54.158 | 55.4834 | SDsa1
21-deoxycortisol | Hormones in Serum |  | 1 | nmol/L |  |  |  |  | SDsa1
25 OH Vitamin D | Hormones in Serum | Biological variation | 50 | nmol/L | 12.442 | 35.2979 | 6.221 | 17.6489 | SDsa1 | 2023-02-09
ACTH | Hormones in Serum |  | 40 | pmol/L |  | 54.6776 |  | 21.871 | SDsa1
Active B12 | Hormones in Serum |  | 90 | pmol/L |  | 14.7776 |  | 13.2999 | SDsa1
Aldosteron | Hormones in Serum | Biological variation | 500 | pmol/L | 42.804 | 25.3606 | 214.02 | 126.803 | TEa | 2021-01-01
AMH | Hormones in Serum | Biological variation | 5 | ug/L | 20.64 | 24.0732 | 1.032 | 1.2037 | SDsa1 | 2022-01-01
Androsteendion | Hormones in Serum | Biological variation | 10 | nmol/L | 23.508 | 20.7448 | 2.3508 | 2.0745 | TEa
Copeptin | Hormones in Serum |  | 3 | pmol/L |  |  |  |  | SDsa1
Corticosterone | Hormones in Serum |  | 50 | nmol/L |  | 15.369 |  | 7.6845 | SDsa1
Cortisol | Hormones in Serum | Biological variation | 500 | nmol/L | 26.286 | 15.6937 | 131.43 | 78.4683 | TEa | 2023-02-09
C-peptide | Hormones in Serum | Biological variation | 1.5 | nmol/L | 20.827 | 18.987 | 0.31241 | 0.2848 | TEa
CTX | Hormones in Serum |  | 700 | ng/L | 24.537 | 34.6171 | 171.759 | 242.3198 | SDsa1 | 2023-02-09
DHEA | Hormones in Serum |  | 20 | nmol/L |  | 27.8065 |  | 5.5613 | SDsa1
DHEAS | Hormones in Serum | Biological variation | 7.5 | umol/L | 19.04 | 23.2763 | 1.428 | 1.7457 | SDsa1 | 2024-02-06
Digoxin | Hormones in Serum |  | 1 | ug/L |  | 26.1508 |  | 0.26151 | SDsa1
Eestrone | Hormones in Serum |  | 400 | pmol/L |  |  |  |  | SDsa1
Ferritin | Hormones in Serum | Biological variation | 50 | ug/L | 13.76 | 13.911 | 6.88 | 6.9555 | SDsa1 | 2021-01-01
Folate | Hormones in Serum | Biological variation | 20 | nmol/L | 19.659 | 25.3187 | 3.9318 | 5.0637 | TEa | 2023-02-09
Free testosteron | Hormones in Serum |  | 500 | pmol/L | 27.25 | 37.3686 | 136.25 | 186.843 | SDsa1 | 2021-01-01
FSH | Hormones in Serum | Biological variation | 15 | [IU]/L | 21.202 | 14.08 | 3.1803 | 2.112 | TEa | 2021-01-01
Growth hormone | Hormones in Serum | Biological variation | 10 | m[IU]/L | 21.6 | 17.4721 | 2.16 | 1.7472 | TEa
IGF-1 | Hormones in Serum | Biological variation | 20 | nmol/L | 14.902 | 28.2615 | 2.9804 | 5.6523 | SDsa1 | 2021-01-01
IGF1-SD | Hormones in Serum |  | 100 |  |  | 50.9377 |  | 50.9377 | SDsa1
Inhibin B | Hormones in Serum |  | 300 | ng/L |  | 23.3683 |  | 70.1048 | SDsa1
Insulin | Hormones in Serum | Biological variation | 33.3333 | m[IU]/L | 31.465 | 22.0355 | 10.4883 | 7.3452 | TEa | 2023-02-09
Leptin | Hormones in Serum |  | 20 | ug/L |  |  |  |  | SDsa1
LH | Hormones in Serum | Biological variation | 10 | [IU]/L | 28.39 | 13.3233 | 2.839 | 1.3323 | TEa | 2021-01-01
Methylmalonic acid | Hormones in Serum |  | 750 | nmol/L |  | 26.7686 |  | 200.7646 | SDsa1
Oestradiol | Hormones in Serum | Biological variation | 500 | pmol/L | 17.337 | 18.224 | 86.685 | 91.1202 | SDsa1 | 2021-01-01
Osteocalcin | Hormones in Serum | Biological variation | 4 | nmol/L | 15.718 |  | 0.62872 |  | SDsa1 | 2023-02-09
P3NP | Hormones in Serum |  | 10 | ug/L |  |  |  |  | SDsa1
PIGF | Hormones in Serum |  | 500 | ng/L |  |  |  |  | SDsa1
PINP | Hormones in Serum |  | 80 | ug/L | 16.258 | 14.3206 | 13.0064 | 11.4565 | SDsa1 | 2024-02-06
Progesteron | Hormones in Serum |  | 15 | nmol/L | 26.212 | 19.5789 | 3.9318 | 2.9368 | SDsa1 | 2023-02-09
Progesteron,17OH | Hormones in Serum | Biological variation | 10 | nmol/L | 35.313 | 22.1844 | 3.5313 | 2.2184 | TEa | 2021-01-01
Prolactin | Hormones in Serum | Biological variation | 0.25 | [IU]/L | 37.374 | 17.9084 | 0.093435 | 0.044771 | TEa | 2022-01-01
PTH | Hormones in Serum | Biological variation | 10 | pmol/L | 18.91 | 23.1093 | 1.891 | 2.3109 | SDsa1 | 2024-02-06
Renin antigen | Hormones in Serum | Biological variation | 60 | m[IU]/L | 37.669 | 33.4983 | 22.6014 | 20.099 | TEa | 2021-01-01
sFlt1 | Hormones in Serum |  | 1000 | ng/L |  |  |  |  | SDsa1
SHBG | Hormones in Serum | Biological variation | 45 | nmol/L | 17.372 | 18.5963 | 7.8174 | 8.3683 | SDsa1 | 2023-02-09
T3 | Hormones in Serum | Biological variation | 3 | nmol/L | 11.605 | 18.0904 | 0.34815 | 0.54271 | SDsa1 | 2021-01-01
T3, free | Hormones in Serum | Biological variation | 10 | pmol/L | 6.547 | 9.9532 | 0.6547 | 0.99532 | SDsa1 | 2023-02-09
T4 | Hormones in Serum | Biological variation | 150 | nmol/L | 8.636 | 15.5289 | 12.954 | 23.2934 | SDsa1 | 2024-02-06
T4, free | Hormones in Serum | Biological variation | 25 | pmol/L | 6.345 | 15.1491 | 1.5862 | 3.7873 | SDsa1 | 2023-02-09
TBG | Hormones in Serum | Biological variation | 20 | mg/L | 0.101 | 29.4753 | 0.0202 | 5.8951 | SDsa1
Testosteron | Hormones in Serum | Biological variation | 20 | nmol/L | 16.508 | 16.2815 | 3.3016 | 3.2563 | TEa | 2022-01-01
TSH | Hormones in Serum | Biological variation | 5 | m[IU]/L | 24.609 | 11.6187 | 1.2305 | 0.58093 | TEa | 2023-02-09
Vit. D 1,25 (OH)2 | Hormones in Serum |  | 100 | pmol/L |  | 41.7011 |  | 41.7011 | SDsa1
Vitamin B12 | Hormones in Serum |  | 500 | pmol/L |  | 13.3234 |  | 66.6168 | SDsa1
3-MT | Hormones in Urine | Biological variation | 400 | nmol/L | 29.8 | 63.6156 | 119.2 | 254.4625 | SDsa1
5-HIAA | Hormones in Urine | Biological variation | 50 | umol/L | 26.476 | 26.9937 | 13.238 | 13.4969 | SDsa1
Adrenaline | Hormones in Urine | Biological variation | 750 | nmol/L | 147.725 | 23.3873 | 1107.9375 | 175.4044 | TEa | 2021-01-01
Aldosteron | Hormones in Urine | Biological variation | 30 | nmol/L | 46.559 |  | 13.9677 |  | TEa
Cortisol | Hormones in Urine |  | 200 | nmol/L |  | 27.3765 |  | 54.7529 | SDsa1
Cortison | Hormones in Urine |  | 120 | nmol/L |  | 35.2656 |  | 42.3187 | SDsa1
Dopamine | Hormones in Urine | Biological variation | 5000 | nmol/L | 29.8 | 22.4578 | 1490 | 1122.8895 | TEa
Free metanefrine | Hormones in Urine |  | 8000 | nmol/L |  |  |  |  | SDsa1
Free normetanephrine | Hormones in Urine | Biological variation | 1000 | nmol/L | 29.8 |  | 298 |  | TEa
Free 3-MT | Hormones in Urine | Biological variation | 400 | nmol/L | 29.8 | 63.6156 | 119.2 | 254.4625 | SDsa1
HVA | Hormones in Urine | Biological variation | 75 | umol/L | 29.8 | 19.4603 | 22.35 | 14.5952 | TEa
Metanefrine | Hormones in Urine | Biological variation | 5000 | nmol/L | 29.8 | 24.8421 | 1490 | 1242.1072 | TEa
Noradrenaline | Hormones in Urine | Biological variation | 500 | nmol/L | 29.8 | 25.0846 | 149 | 125.4231 | TEa
Normetanephrine | Hormones in Urine | Biological variation | 1000 | nmol/L | 29.8 | 38.0476 | 298 | 380.476 | SDsa1
VMA | Hormones in Urine | Biological variation | 50 | umol/L | 31.31 | 19.1684 | 15.655 | 9.5842 | TEa
AS aangetoond tegen bofvirus? | Immunity serology |  |  | Qualit.
AS aangetoond tegen SARS-CoV-2? | Immunity serology |  |  | Qualit.
Beschermd tegen hepatitis A? | Immunity serology |  |  | Qualit.
Beschermd tegen hepatitis B? | Immunity serology |  |  | Qualit.
Beschermd tegen mazelen? | Immunity serology |  |  | Qualit.
Beschermd tegen rubella? | Immunity serology |  |  | Qualit.
Beschermd tegen waterpokken? | Immunity serology |  |  | Qualit.
Bof antibodies | Immunity serology |  |  | {titer}
Bof IgG | Immunity serology |  |  | IZS **
Hepatitis A IgG | Immunity serology |  |  | {ratio}
Hepatitis A IgT | Immunity serology |  |  | {ratio}
Hepatitis B Anti-HBs | Immunity serology |  |  | [IU]/L
Measels antibodies | Immunity serology |  |  | IZS **
Measels IgG | Immunity serology |  |  | IZS **
Rubella antibodies | Immunity serology |  |  | IZS **
Rubella IgG | Immunity serology |  |  | [IU]/mL
SARS-CoV IgG | Immunity serology |  |  | k[arb'U]/L
Varicella zoster antibodies | Immunity serology |  |  | IZS **
Varicella zoster IgG | Immunity serology |  |  | IZS **
Ciclosporin | Immunosuppressant microsampling | Based on EMA/FDA | 150 | ug/L | 26.5 |  | 39.75 |  | SDsa1
Creatinine | Immunosuppressant microsampling | Biological variation | 100 | umol/l | 7.46 |  | 7.46 |  | SDsa1 | 2022-01-01
Everolimus | Immunosuppressant microsampling | Based on EMA/FDA | 5 | ug/L | 26.5 |  | 1.325 |  | SDsa1
Hematocrit | Immunosuppressant microsampling | Biological variation | 0.4 | L/L | 3.875 |  | 0.0155 |  | SDsa1 | 2024-01-30
Mycophenolic acid | Immunosuppressant microsampling | Based on EMA/FDA | 2.5 | mg/L | 26.5 |  | 0.6625 |  | SDsa1
Sirolimus | Immunosuppressant microsampling | Based on EMA/FDA | 5 | ug/L | 26.5 |  | 1.325 |  | SDsa1
Tacrolimus | Immunosuppressant microsampling | Based on EMA/FDA | 5 | ug/L | 26.5 |  | 1.325 |  | SDsa1
CRP | Inflammation markers | Biological variation | 10 | mg/L | 50.704 | 17.0922 | 5.0704 | 1.7092 | TEa | 2021-01-01
Procalcitonine | Inflammation markers |  | 2.5 | ug/L |  | 13.2009 |  | 0.33002 | SDsa1
Blastocystis spp. PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  |  |  |  | SDsa1
Cryptosporidium spp. PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  | 30.607 |  | 6.1214 | SDsa1
Dientamoeba fragilis PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  | 18.7685 |  | 3.7537 | SDsa1
Entamoeba dispar PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  |  |  |  | SDsa1
Entamoeba histolytica PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  | 31.0958 |  | 6.2192 | SDsa1
Giardia lamblia PCR | Intestinal protozoa Molecular Diagnostics |  | 20 | {Cq} |  | 39.3096 |  | 7.8619 | SDsa1
Ectoparasite | Intestinal-, tissue and ecto parasites |  |  | Qualit.
Parasite intestinal/tissue | Intestinal-, tissue and ecto parasites |  |  | Qualit.
ALAT | KILO | Biological variation | 50 | U/L | 16.08 | 14.4945 | 8.04 | 7.2472 | TEa | 2021-01-01
Alk. Phosphatase | KILO | Biological variation | 150 | U/L | 10.541 | 13.1528 | 15.8115 | 19.7293 | SDsa1 | 2022-01-01
Amylase | KILO | Biological variation | 250 | U/L | 13.173 | 9.9744 | 32.9325 | 24.9361 | TEa | 2021-01-01
ASAT | KILO | Biological variation | 50 | U/L | 13.511 | 13.3811 | 6.7555 | 6.6906 | TEa | 2021-01-01
Bilirubin | KILO | Biological variation | 20 | umol/L | 24.565 | 12.9837 | 4.913 | 2.5967 | TEa | 2024-01-18
Calcium | KILO | Biological variation | 2.5 | mmol/L | 2.296 | 6.0808 | 0.0574 | 0.15202 | SDsa1 | 2022-01-01
Chloride | KILO | Biological variation | 100 | mmol/L | 1.333 | 4.9584 | 1.333 | 4.9584 | SDsa1 | 2021-01-01
Cholesterol | KILO | Biological variation | 5 | mmol/L | 8.658 | 7.4177 | 0.4329 | 0.37088 | TEa | 2021-01-01
CK | KILO | Biological variation | 150 | U/L | 22.565 | 9.1422 | 33.8475 | 13.7133 | TEa | 2021-01-01
Creatinine | KILO | Biological variation | 100 | umol/L | 7.413 | 8.3015 | 7.413 | 8.3015 | SDsa1 | 2022-01-01
Gamma-GT | KILO | Biological variation | 50 | U/L | 18.257 | 9.8993 | 9.1285 | 4.9496 | TEa | 2021-01-01
Glucose | KILO | Biological variation | 10 | mmol/L | 10 | 6.9615 | 1 | 0.69615 | TEa | 2021-01-01
Inorg. Phosphate | KILO | Biological variation | 1 | mmol/L | 9.745 | 8.495 | 0.09745 | 0.08495 | TEa
Iron | KILO | Biological variation | 30 | umol/L | 26.668 | 7.194 | 8.0004 | 2.1582 | TEa
LD | KILO | Biological variation | 500 | U/L | 7.698 | 12.8753 | 38.49 | 64.3764 | SDsa1 | 2021-01-01
Potassium | KILO | Biological variation | 4 | mmol/L | 4.85 | 4.357 | 0.194 | 0.17428 | TEa | 2021-01-01
Sodium | KILO | Biological variation | 140 | mmol/L | 0.692 | 2.876 | 0.9688 | 4.0264 | SDsa1 | 2021-01-01
Total Protein | KILO | Biological variation | 70 | g/L | 3.466 | 5.8259 | 2.4262 | 4.0781 | SDsa1 | 2021-01-01
Triglycerides | KILO | Biological variation | 2 | mmol/L | 27.037 | 10.3797 | 0.54074 | 0.20759 | TEa | 2021-01-01
Urate | KILO | Biological variation | 500 | umol/L | 12.82 | 7.108 | 64.1 | 35.5398 | TEa | 2022-01-01
Urea | KILO | Biological variation | 10 | mmol/L | 17.592 | 8.7924 | 1.7592 | 0.87924 | TEa
Abberating population | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
B-lymphocytes | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 1 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 10 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 103 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 109 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 117 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 11a | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 11b | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 11c | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 123 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 13 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 133 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 135 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 138/Vs38c | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 14 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 15 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 16 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 18 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 19 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 2 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 20 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 200 | Leukemia/Lymphoma typing |  | 100 | % | 8.90 |  | 8.9 |  | TEa
CD 21 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 22 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 23 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 235a | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 24 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 25 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 3 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 30 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 33 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 34 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 36 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 37 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 38 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 4 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 41 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 42 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 43 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 45 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 45RA | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 45RO | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 5 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 56 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 57 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 61 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 62P | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 64 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 65s | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 66 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 66e | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 7 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 71 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 79a | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 79b | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 8 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD 9 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyCD 3 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyCD22 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyCD79a | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-A | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-D | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-G | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-Kappa | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-Lambda | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CyIg-M | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
FMC-7 | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
HLADR | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
Monocytair | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
MPO | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
Myeloïd | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
NK-lymphocytes | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-A | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-D | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-G | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-Kappa | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-Lambda | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
sIg-M | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
TCR-alpha-beta | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
TCRgd | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
TdT | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
T-lymphocytes | Leukemia/Lymphoma typing |  | 100 | % | 8.9 |  | 8.9 |  | TEa
Apo A1 lipoproteïn | Lipids | Biological variation | 1.4 | g/L | 7.563 | 13.8633 | 0.10588 | 0.19409 | SDsa1 | 2022-01-01
Apo B lipoproteïn | Lipids | Biological variation | 1 | g/L | 11.483 | 13.6529 | 0.11483 | 0.13653 | SDsa1 | 2021-01-01
Cholesterol | Lipids | Biological variation | 5 | mmol/L | 8.658 | 7.0973 | 0.4329 | 0.35486 | TEa | 2021-01-01
Cholesterol ratio | Lipids | Biological variation | 3.5 |  | 12 | 8.7331 | 0.42 | 0.30566 | TEa
HDL-Cholesterol | Lipids | Biological variation | 1 | mmol/L | 11.079 | 9.7183 | 0.11079 | 0.097183 | TEa | 2021-01-01
hsCRP | Lipids | Biological variation | 10 | mg/L | 72.908 |  | 7.2908 |  | SDsa1
LDL-Cholesterol | Lipids | Biological variation | 3 | mmol/L | 13.694 | 11.0084 | 0.41082 | 0.33025 | TEa | 2021-01-01
Lp(a) | Lipids | Biological variation | 50 | nmol/L | 9.568 |  | 4.784 |  | SDsa1 | 2022-01-01
non-HDL cholesterol | Lipids | Biological variation | 3.6 | mmol/L | 9.965 | 8.683 | 0.35874 | 0.31259 | TEa | 2023-02-14
Triglycerides | Lipids | Biological variation | 2 | mmol/L | 27.037 | 9.2532 | 0.54074 | 0.18506 | TEa | 2021-01-01
Albumine liquor | Liquor protein |  | 200 | mg/L |  | 17.1751 |  | 34.3502 | SDsa1
Albumine serum | Liquor protein | Biological variation | 40 | g/L | 3.532 | 13.6263 | 1.4128 | 5.4505 | SDsa1 | 2024-01-15
IgG index | Liquor protein |  | 1.2 | {ratio} |  | 25.0555 |  | 0.30067 | SDsa1
IgG liquor | Liquor protein |  | 30 | mg/L |  | 14.8149 |  | 4.4445 | SDsa1
IgG serum | Liquor protein | Biological variation | 13 | g/L | 7.251 | 8.5286 | 0.94263 | 1.1087 | SDsa1 | 2021-01-01
intrathecale IgG fractie | Liquor protein |  | 40 | % |  | 47.663 |  | 19.0652 | SDsa1
intrathecale IgG fractie (kwal.) | Liquor protein |  |  | Qualit.
intrathecale IgG production | Liquor protein |  | 60 | mg/L |  |  |  |  | SDsa1
Liquor bands | Liquor protein |  |  | Qualit.
Q Albumin | Liquor protein |  | 4.5 | {ratio} |  | 18.8564 |  | 0.84854 | SDsa1
Q IgG | Liquor protein |  | 3 | {ratio} |  | 19.5119 |  | 0.58536 | SDsa1
Serum bands | Liquor protein |  |  | Qualit.
Total Protein liquor | Liquor protein |  | 400 | mg/L |  | 13.6351 |  | 54.5405 | SDsa1
017 (T. pal) | Lues / Lyme serology |  |  | Qualit.
B. burgdorferi IgG aanwezig? | Lues / Lyme serology |  |  | Qualit.
B. burgdorferi IgG blot? | Lues / Lyme serology |  |  | Qualit.
B. burgdorferi IgM aanwezig? | Lues / Lyme serology |  |  | Qualit.
B. burgdorferi IgM blot? | Lues / Lyme serology |  |  | Qualit.
Borrelia burgdorferi IgG | Lues / Lyme serology |  |  | IZS **
Borrelia burgdorferi IgM | Lues / Lyme serology |  |  | IZS **
Borrelia IgTotal | Lues / Lyme serology |  |  | IZS **
CL1 (T. pal) | Lues / Lyme serology |  |  | Qualit.
CL2 (T. pal) | Lues / Lyme serology |  |  | Qualit.
CL3 (T. pal) | Lues / Lyme serology |  |  | Qualit.
DbpA (IgG blot) | Lues / Lyme serology |  |  | Qualit.
DbpA (IgM blot) | Lues / Lyme serology |  |  | Qualit.
iv 1 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
iv 2 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
iv 3 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
iv 4 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
Lues infectie? | Lues / Lyme serology |  |  | Qualit.
Lyme behandeling? | Lues / Lyme serology |  |  | Qualit.
Non-treponemale antistoffen aanwezig? | Lues / Lyme serology |  |  | Qualit.
Osp17 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
OspA (IgG blot) | Lues / Lyme serology |  |  | Qualit.
OspA (IgM blot) | Lues / Lyme serology |  |  | Qualit.
OspC (IgG blot) | Lues / Lyme serology |  |  | Qualit.
OspC (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p100 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p100 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p14 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p15 (T. pal) | Lues / Lyme serology |  |  | Qualit.
p17 (T. pal) | Lues / Lyme serology |  |  | Qualit.
p18 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p18 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p30 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p39 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p39 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p41 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p41 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p43 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p44.5 (T. pal) | Lues / Lyme serology |  |  | Qualit.
p45 (T. pal) | Lues / Lyme serology |  |  | Qualit.
p47 (T. pal) | Lues / Lyme serology |  |  | Qualit.
p58 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
p58 (IgM blot) | Lues / Lyme serology |  |  | Qualit.
p83 (IgG blot) | Lues / Lyme serology |  |  | Qualit.
T. pallidum IgG blot? | Lues / Lyme serology |  |  | Qualit.
T. pallidum FTA. | Lues / Lyme serology |  |  | {titer}
T. pallidum IgM FTA | Lues / Lyme serology |  |  | Qualit.
T. pallidum screening | Lues / Lyme serology |  |  | {ratio}
T. pallidum TPHA/TPPA | Lues / Lyme serology |  |  | {titer}
TmpA (T. pal) | Lues / Lyme serology |  |  | Qualit.
Tp15 (T. pal) | Lues / Lyme serology |  |  | Qualit.
TP17 (T. pal) | Lues / Lyme serology |  |  | Qualit.
Tp257 (T. pal) | Lues / Lyme serology |  |  | Qualit.
Tp453 (T. pal) | Lues / Lyme serology |  |  | Qualit.
TpN15 (T. pal) | Lues / Lyme serology |  |  | Qualit.
TpN17 (T. pal) | Lues / Lyme serology |  |  | Qualit.
TpN47 (T. pal) | Lues / Lyme serology |  |  | Qualit.
TPp47 (T. pal) | Lues / Lyme serology |  |  | Qualit.
Treponemale antistoffen aanwezig? | Lues / Lyme serology |  |  | Qualit.
VDRL/RPR | Lues / Lyme serology |  |  | {titer}
vdrl1 (T. pal) | Lues / Lyme serology |  |  | Qualit.
vdrl2 (T. pal) | Lues / Lyme serology |  |  | Qualit.
vdrl3 (T. pal) | Lues / Lyme serology |  |  | Qualit.
vdrl4 (T. pal) | Lues / Lyme serology |  |  | Qualit.
vdrl5 (T. pal) | Lues / Lyme serology |  |  | Qualit.
VlsE (IgG blot) | Lues / Lyme serology |  |  | Qualit.
VLsE (IgM blot) | Lues / Lyme serology |  |  | Qualit.
% pos macrophagen | Lymphocyte subsets BAL |  | 40 | % | 40 |  | 16 |  | TEa
CD103 | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
CD103,CD4/CD4 | Lymphocyte subsets BAL |  | 0.25 | {ratio} | 25 |  | 0.0625 |  | TEa
CD19+ | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
CD3+ | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
CD3+CD4+ | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
CD3+CD8+ | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
CD3-CD16+CD56+ | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Eosinophils | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Golde score | Lymphocyte subsets BAL |  | 2 | N/A | 50 |  | 1 |  | TEa
Incubation temperature | Lymphocyte subsets BAL |  |  | Qualit.
Lymphocytes | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
Lymphocytes on flow | Lymphocyte subsets BAL |  | 20 | % | 25 |  | 5 |  | TEa
Macrophages | Lymphocyte subsets BAL |  | 25 | % | 25 |  | 6.25 |  | TEa
Mast cells | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Neutrophils | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Number of cells | Lymphocyte subsets BAL |  | 50 | 10*4/mL | 25 |  | 12.5 |  | TEa
Other | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Plasma cells | Lymphocyte subsets BAL |  | 10 | % | 25 |  | 2.5 |  | TEa
Alfa thal. - alfa 3.7 | Molecular Biology various |  |  | Qualit.
Alfa thal. - alfa 4.2 | Molecular Biology various |  |  | Qualit.
Alfa thal. --20.5 | Molecular Biology various |  |  | Qualit.
Alfa thal. --FIL | Molecular Biology various |  |  | Qualit.
Alfa thal. --MED | Molecular Biology various |  |  | Qualit.
Alfa thal. --SEA | Molecular Biology various |  |  | Qualit.
Alfa thal. --THAI | Molecular Biology various |  |  | Qualit.
Alfa thal. --Triplicatie | Molecular Biology various |  |  | Qualit.
CF-delta 508 | Molecular Biology various |  |  | Qualit.
HFE C282Y | Molecular Biology various |  |  | Qualit.
HFE H63D | Molecular Biology various |  |  | Qualit.
HLA-B27 | Molecular Biology various |  |  | Qualit.
HLA-DQ2.2 | Molecular Biology various |  |  | Qualit.
HLA-DQ2.5 | Molecular Biology various |  |  | Qualit.
HLA-DQ8 | Molecular Biology various |  |  | Qualit.
LCT -13910C>T | Molecular Biology various |  |  | Qualit.
CALR Mutation | MPN molecular diagnostics |  |  | Qualit.
CALR type of mutation | MPN molecular diagnostics |  |  | Qualit.
CALR VAF | MPN molecular diagnostics |  |  | % |  |  |  |  | SDsa1
JAK2 exon 12 mutation | MPN molecular diagnostics |  |  | Qualit.
JAK2 V617F | MPN molecular diagnostics |  | 10 | % |  | 67.1864 |  | 6.7186 | SDsa1
MPL mutation | MPN molecular diagnostics |  |  | Qualit.
Albumin spectrum | M-protein |  | 40 | g/L |  | 11.7307 |  | 4.6923 | SDsa1
Alpha-1-globulin | M-protein | Biological variation | 2 | g/L | 15.733 | 30.0305 | 0.31466 | 0.60061 | SDsa1
Alpha-2-globulin | M-protein | Biological variation | 7 | g/L | 12.585 | 35.0696 | 0.88095 | 2.4549 | SDsa1
Beta-globulin | M-protein | Biological variation | 7.3 | g/L | 11.731 | 35.7076 | 0.85636 | 2.6067 | SDsa1
Cryoglobulin | M-protein |  | 70 | mg/L |  |  |  |  | SDsa1
Free kappa | M-protein |  |  | g/L
Free lambda | M-protein |  |  | g/L
Gamma-globulin | M-protein | Biological variation | 8.3 | g/L | 16.818 | 25.562 | 1.3959 | 2.1216 | SDsa1
IgA | M-protein | Biological variation | 3 | g/L | 9.782 | 9.8083 | 0.29346 | 0.29425 | SDsa1 | 2021-01-01
IgA (light chain ?) | M-protein |  |  | g/L
IgA heavy chain | M-protein |  |  | g/L
IgA kappa | M-protein |  |  | g/L
IgA kappa 2nb band | M-protein |  |  | g/L
IgA kappa 3rd band | M-protein |  |  | g/L
IgA lambda | M-protein |  |  | g/L
IgA lambda 2nd band | M-protein |  |  | g/L
IgA lambda 3rd band | M-protein |  |  | g/L
IgD | M-protein |  | 63.45 | mg/L |  |  |  |  | SDsa1
IgD (light chain ?) | M-protein |  |  | g/L
IgD kappa | M-protein |  |  | g/L
IgD lambda | M-protein |  |  | g/L
IgG | M-protein | Biological variation | 13 | g/L | 7.251 | 9.5192 | 0.94263 | 1.2375 | SDsa1 | 2021-01-01
IgG (light chain ?) | M-protein |  |  | g/L
IgG heavy chain | M-protein |  |  | g/L
IgG kappa | M-protein |  |  | g/L
IgG kappa 2nb band | M-protein |  |  | g/L
IgG kappa 3rd band | M-protein |  |  | g/L
IgG lambda | M-protein |  |  | g/L
IgG lambda 2nd band | M-protein |  |  | g/L
IgG lambda 3rd band | M-protein |  |  | g/L
IgM | M-protein | Biological variation | 1.2 | g/L | 17.082 | 12.4746 | 0.20498 | 0.1497 | SDsa1 | 2021-01-01
IgM (light chain ?) | M-protein |  |  | g/L
IgM heavy chain | M-protein |  |  | g/L
IgM kappa | M-protein |  |  | g/L
IgM kappa 2nd band | M-protein |  |  | g/L
IgM kappa 3rd band | M-protein |  |  | g/L
IgM lamba 2nd band | M-protein |  |  | g/L
IgM lamba 3rd band | M-protein |  |  | g/L
IgM lambda | M-protein |  |  | g/L
Ratio (kappa/lambda) | M-protein |  | 40 | {ratio} |  | 83.0944 |  | 33.2377 | SDsa1
sVLK kappa | M-protein |  | 400 | mg/L |  | 54.1411 |  | 216.5644 | SDsa1
sVLK lambda | M-protein |  | 100 | mg/L |  | 34.2979 |  | 34.2979 | SDsa1
Total Protein | M-protein | Biological variation | 70 | g/L | 3.466 | 6.5474 | 2.4262 | 4.5832 | SDsa1 | 2021-01-01
Total Protein (urin) | M-protein | Biological variation | 70 | g/L | 39.959 |  | 27.9713 |  | TEa
Determination | Mycobacteriology |  |  | Qualit.
Determination | Mycology |  |  | Qualit.
Albumin | Neonatal Bilirubin | Biological variation | 20 | g/L | 3.532 | 13.1762 | 0.7064 | 2.6352 | SDsa1 | 2024-01-25
Bilirubin | Neonatal Bilirubin | Clinical decision | 350 | umol/L | 5 | 8.8889 | 17.5 | 31.111 | SDsa1
Creatinine | Neonatal Bilirubin | Biological variation | 25 | umol/L | 7.413 | 26.3113 | 1.8533 | 6.5778 | SDsa1 | 2022-01-01
Glucose (low) | Neonatal Bilirubin | Biological variation | 3 | mmol/L | 6.505 | 7.494 | 0.22482 |  | SDsa1 | 2021-01-01
Hb in feces | Occult blood in feces |  | 250 | ug/L |  | 54.0125 |  | 135.0312 | SDsa1
Hb in feces (qual.) | Occult blood in feces |  |  | qualit.
Cabozantinib | Oral oncolytics | Based on EMA/FDA | 100 | ug/L | 26.5 |  | 26.5 |  | SDsa1
Desethylsunitinib | Oral oncolytics | Based on EMA/FDA | 5 | ug/L | 26.5 |  | 1.325 |  | SDsa1
Imatinib | Oral oncolytics | Based on EMA/FDA | 5 | mg/L | 26.5 | 33.3586 | 1.325 | 1.6679 | SDsa1
Olaparib | Oral oncolytics | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | SDsa1
Pazopanib | Oral oncolytics | Based on EMA/FDA | 25 | mg/L | 26.5 | 11.4672 | 6.625 | 2.8668 | TEa
Sunitinib | Oral oncolytics | Based on EMA/FDA | 25 | ug/L | 26.5 |  |  |  | SDsa1
c.1236G>A | Pharmacogenetics |  |  | Qualit.
c.1679T>G | Pharmacogenetics |  |  | Qualit.
c.2846A>T | Pharmacogenetics |  |  | Qualit.
CYP2C19*2 | Pharmacogenetics |  |  | Qualit.
CYP2D6*3 | Pharmacogenetics |  |  | Qualit.
CYP2D6*4 | Pharmacogenetics |  |  | Qualit.
CYP2D6*41 | Pharmacogenetics |  |  | Qualit.
CYP2D6*5 (deletie) | Pharmacogenetics |  |  | Qualit.
CYP2D6*6 | Pharmacogenetics |  |  | Qualit.
CYP2D6*duplicatie | Pharmacogenetics |  |  | Qualit.
DPYD*2A (IVS14+1 G>A) | Pharmacogenetics |  |  | Qualit.
Albumin | Plasma proteins | Biological variation | 50 | g/L | 3.532 | 11.5865 | 1.766 | 5.7932 | SDsa1 | 2024-01-22
Alpha-1-antitrypsin | Plasma proteins | Biological variation | 1.7 | g/L | 6.201 | 12.3986 | 0.10542 | 0.21078 | SDsa1 | 2024-01-22
Beta-2-microglobulin | Plasma proteins | Biological variation | 3 | mg/L | 6.435 | 12.4848 | 0.19305 | 0.37454 | SDsa1 | 2021-01-01
C3 | Plasma proteins | Biological variation | 1.8 | g/L | 7.765 | 12.2721 | 0.13977 | 0.2209 | SDsa1 | 2021-01-01
C3d | Plasma proteins |  | 1 | mg/L |  |  |  |  | SDsa1
C4 | Plasma proteins | Biological variation | 0.35 | g/L | 12.056 | 12.2576 | 0.042196 | 0.042902 | SDsa1 | 2021-01-01
Ceruloplasmin | Plasma proteins | Biological variation | 450 | mg/L | 8.035 | 13.671 | 36.1575 | 61.5195 | SDsa1 | 2021-01-01
CRP | Plasma proteins | Biological variation | 10 | mg/L | 50.704 | 12.5594 | 5.0704 | 1.2559 | TEa | 2021-01-01
Ferritin | Plasma proteins | Biological variation | 200 | ug/L | 13.76 | 13.7477 | 27.52 | 27.4953 | TEa | 2021-01-01
Haptoglobin | Plasma proteins | Biological variation | 2 | g/L | 17.079 | 7.9376 | 0.34158 | 0.15875 | TEa | 2021-01-01
IgA | Plasma proteins | Biological variation | 3 | g/L | 9.782 | 9.8083 | 0.29346 | 0.29425 | SDsa1 | 2021-01-01
IgA1 | Plasma proteins |  | 2 | g/L |  |  |  |  | SDsa1
IgA2 | Plasma proteins |  | 0.5 | g/L |  |  |  |  | SDsa1
IgD | Plasma proteins |  | 63.45 | mg/L |  |  |  |  | SDsa1
IgE | Plasma proteins |  | 150 | [IU]/mL |  | 24.1394 |  | 36.209 | SDsa1
IgG | Plasma proteins | Biological variation | 13 | g/L | 7.251 | 9.5192 | 0.94263 | 1.2375 | SDsa1 | 2021-01-01
IgG1 | Plasma proteins |  | 8 | g/L |  | 10.2771 |  | 0.82217 | SDsa1
IgG2 | Plasma proteins |  | 4 | g/L |  | 12.5055 |  | 0.50022 | SDsa1
IgG3 | Plasma proteins |  | 0.6 | g/L |  | 17.5987 |  | 0.10559 | SDsa1
IgG4 | Plasma proteins |  | 0.7 | g/L |  | 13.7525 |  | 0.096267 | SDsa1
IgM | Plasma proteins | Biological variation | 1.2 | g/L | 17.082 | 12.4746 | 0.20498 | 0.1497 | TEa | 2021-01-01
Iron saturation % | Plasma proteins |  | 30 | % |  | 14.7382 |  | 4.4215 | SDsa1
Prealbumin | Plasma proteins | Biological variation | 300 | mg/L | 14.49 | 13.7607 | 43.47 | 41.2822 | TEa
Transferrin | Plasma proteins | Biological variation | 3 | g/L | 6.827 | 8.5814 | 0.20481 | 0.25744 | SDsa1 | 2021-01-01
Zure alfa-1-glyco | Plasma proteins | Biological variation | 1 | g/L | 12.408 |  | 0.12408 |  | SDsa1 | 2022-01-01
anti-GPIa/IIa | Platelets / leukocyte serology |  |  | Qualit.
anti-GPIb/IX | Platelets / leukocyte serology |  |  | Qualit.
anti-GPIIb/IIIa | Platelets / leukocyte serology |  |  | Qualit.
anti-GPIV | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-15a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-15b | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-1a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-1b | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-2a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-2b | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-3a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-3b | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-4a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-4b | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-5a | Platelets / leukocyte serology |  |  | Qualit.
anti-HPA-5b | Platelets / leukocyte serology |  |  | Qualit.
Aspecific | Platelets / leukocyte serology |  |  | Qualit.
HLA-Class-I antibodies | Platelets / leukocyte serology |  |  | Qualit.
HPA-15a | Platelets / leukocyte serology |  |  | Qualit.
HPA-15b | Platelets / leukocyte serology |  |  | Qualit.
HPA-1a | Platelets / leukocyte serology |  |  | Qualit.
HPA-1b | Platelets / leukocyte serology |  |  | Qualit.
HPA-2a | Platelets / leukocyte serology |  |  | Qualit.
HPA-2b | Platelets / leukocyte serology |  |  | Qualit.
HPA-3a | Platelets / leukocyte serology |  |  | Qualit.
HPA-3b | Platelets / leukocyte serology |  |  | Qualit.
HPA-4a | Platelets / leukocyte serology |  |  | Qualit.
HPA-4b | Platelets / leukocyte serology |  |  | Qualit.
HPA-5a | Platelets / leukocyte serology |  |  | Qualit.
HPA-5b | Platelets / leukocyte serology |  |  | Qualit.
Platelet-antibodies (incl. HLA-ab) | Platelets / leukocyte serology |  |  | Qualit.
CD14 PNH clone (monocytes) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD157 PNH kloon (monocytes) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD157 PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD16 PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD24 PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD55 PNH clone (monocytes) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD55 PNH clone total | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD55 PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD55 Type II cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD55 Type III cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD58 PNH clone total | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD58 Type II cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD58 Type III cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD59 PNH clone (monocytes) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD59 PNH clone total | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD59 PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD59 Type II cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
CD59 Type III cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
FLAER PNH kloon (monocytes) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
FLAER PNH kloon (neutrophils) | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
Monocytes | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
Monocytes (within WBC) | PNH diagnostics |  | 6.5 | % |  |  |  |  | TEa
Neutrophils | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
neutrophils (within WBC) | PNH diagnostics |  | 64.5 | % |  |  |  |  | TEa
Red blood cells | PNH diagnostics |  | 100 | % | 8.9 |  | 8.9 |  | TEa
White blood cells | PNH diagnostics |  | 6.5 | 10*9/L |  |  |  |  | TEa
3-OH-butyric acid | POCT in whole blood |  | 1 | mmol/L |  | 48.0021 |  | 0.48002 | SDsa1
Creatinine | POCT in whole blood | Biological variation | 100 | umol/L | 7.413 |  | 7.413 |  | SDsa1 | 2022-01-01
CRP. | POCT in whole blood | Biological variation | 10 | mg/L | 50.704 | 26.0742 | 5.0704 | 2.6074 | TEa | 2021-01-01
D-Dimer | POCT in whole blood | Biological variation | 1300 | ng/mL | 31.653 | 32.7836 | 411.489 | 426.1868 | SDsa1 | 2024-01-18
Glucose. | POCT in whole blood | Biological variation | 10 | mmol/L | 6.505 | 9.7019 | 0.6505 | 0.97019 | SDsa1 | 2021-01-01
Hemoglobin. | POCT in whole blood | Biological variation | 7 | mmol/L | 3.895 | 5.5963 | 0.27265 | 0.39174 | SDsa1 | 2024-01-30
hs-Troponin I. | POCT in whole blood | Biological variation | 1 | ng/L | 19.363 |  | 0.19363 |  | SDsa1 | 2023-01-31
hs-Troponin T. | POCT in whole blood | Biological variation | 0.014 | ng/L | 30 |  | 0.0042 |  | SDsa1 | 2023-01-31
Lactate | POCT in whole blood | Biological variation | 2 | mmol/L | 30.419 |  | 0.60838 |  | TE
NT-proBNP. | POCT in whole blood | Biological variation | 200 | ng/L | 12.967 |  | 25.934 |  | SDsa1
Potassium | POCT in whole blood | Biological variation | 4 | mmol/L | 4.85 |  | 0.194 |  | SDsa1 | 2021-01-01
Sodium | POCT in whole blood | Biological variation | 140 | mmol/L | 0.692 | 4.1234 | 0.9688 | 5.7728 | SDsa1 | 2021-01-01
Troponine T | POCT in whole blood | Biological variation | 0.1 | ug/L | 48.919 |  | 0.048919 |  | TEa
Coproporphyrin I | Porphyrins in urine |  | 15 | nmol/L |  | 41.8228 |  | 6.2734 | SDsa1
Coproporphyrin I (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Coproprophyrin III | Porphyrins in urine |  | 15 | nmol/L |  | 48.221 |  | 7.2331 | SDsa1
Coproprophyrin III (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Creatinine | Porphyrins in urine | Biological variation | 6 | mmol/L | 28.374 | 9.6052 | 1.7024 | 0.57631 | SDsa1
D-aminolevulinic acid | Porphyrins in urine |  | 50 | umol/L |  | 39.7287 |  | 19.8643 | SDsa1
D-aminolevulinic acid (per creatinine) | Porphyrins in urine |  | 1 | umol/mmol{creat} |  |  |  |  | SDsa1
Heptaporphyrin | Porphyrins in urine |  | 7 | nmol/L |  | 52.138 |  | 3.6497 | SDsa1
Heptaporphyrin (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Hexaporphyrin | Porphyrins in urine |  | 1.5 | nmol/L |  | 64.8556 |  | 0.97283 | SDsa1
Hexaporphyrin (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Pentaporphyrin | Porphyrins in urine |  | 5 | nmol/L |  | 71.8242 |  | 3.5912 | SDsa1
Pentaporphyrin (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Porphobilinogen | Porphyrins in urine | Biological variation | 60 | umol/L | 22.864 | 41.805 | 13.7184 | 25.083 | SDsa1
Porphobilinogen (per creatinine) | Porphyrins in urine |  | 1 | umol/mmol{creat} |  |  |  |  | SDsa1
Total Coproporfyrin | Porphyrins in urine |  | 15 | nmol/L |  |  |  |  | SDsa1
Total Coproporfyrin (per Creatinin) | Porphyrins in urine |  | 50 | umol/mmol{creat} |  |  |  |  | SDsa1
Total Porphyrin | Porphyrins in urine | Biological variation | 700 | nmol/L | 43 | 41.7646 | 301 | 292.3523 | TE
Total Porphyrin (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Total Uroporphyrin | Porphyrins in urine |  | 250 | nmol/L |  | 44.2161 |  | 110.5404 | SDsa1
Total Uroporphyrin (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Uroporphyrin I | Porphyrins in urine |  | 100 | nmol/L |  | 39.3999 |  | 39.3999 | SDsa1
Uroporphyrin I (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Uroporphyrin III | Porphyrins in urine |  | 100 | nmol/L |  |  |  |  | SDsa1
Uroporphyrin III (per creatinine) | Porphyrins in urine |  | 1 | nmol/mmol{creat} |  |  |  |  | SDsa1
Amitriptyline | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 16.4009 | 53 | 32.8018 | TEa
Aripiprazol | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 25.3339 | 53 | 50.6677 | TEa
Citalopram | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 16.8975 | 26.5 | 16.8975 | TEa
Clomipramine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 25.6565 | 26.5 | 25.6565 | TEa
Clopentixol | Psychotherapeutic drugs | Based on EMA/FDA | 10 | ug/L | 26.5 | 28.5208 | 2.65 | 2.8521 | SDsa1
Clozapine | Psychotherapeutic drugs | Based on EMA/FDA | 500 | ug/L | 26.5 | 18.1797 | 132.5 | 90.8984 | TEa
Dehydro-aripiprazole | Psychotherapeutic drugs | Based on EMA/FDA | 62.5 | ug/L | 62.5 |  | 39.0625 |  | TEa
Desipramine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 19.6451 | 26.5 | 19.6451 | TEa
Desmethylcitalopram | Psychotherapeutic drugs | Based on EMA/FDA | 50 | ug/L | 26.5 | 21.1962 | 13.25 | 10.5981 | TEa
Desmethylclomipramine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 20.2328 | 53 | 40.4656 | TEa
Desmethylclozapine | Psychotherapeutic drugs | Based on EMA/FDA | 275 | ug/L | 26.5 |  | 72.875 |  | SDsa1
Desmethylfluoxetine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 29.4867 | 53 | 58.9734 | SDsa1
Desmethylmirtazapine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 22.3629 | 26.5 | 22.3629 | TEa
Desmethylsertraline | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 |  | 53 |  | TEa
Desmethylvenlafaxine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 21.6736 | 53 | 43.3473 | TEa
E-10-OH-Nortriptyline | Psychotherapeutic drugs | Based on EMA/FDA | 137.5 | ug/L | 26.5 |  | 36.4375 |  | SDsa1
Fluoxetine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 23.3434 | 53 | 46.6868 | TEa
Fluvoxamine | Psychotherapeutic drugs | Based on EMA/FDA | 150 | ug/L | 26.5 | 24.3814 | 39.75 | 36.5721 | TEa
Haloperidol | Psychotherapeutic drugs | Based on EMA/FDA | 10 | ug/L | 26.5 | 21.87 | 2.65 | 2.187 | TEa
Hydroxyrisperidon | Psychotherapeutic drugs | Based on EMA/FDA | 40 | ug/L | 26.5 | 24.9117 | 10.6 | 9.9647 | TEa
Imipramine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 19.105 | 26.5 | 19.105 | TEa
Lithium | Psychotherapeutic drugs | Based on EMA/FDA | 1 | mmol/L | 26.5 | 11.8284 | 0.265 | 0.11828 | TEa
Mirtazapine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 16.8254 | 26.5 | 16.8254 | TEa
Nortriptyline | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 20.2831 | 26.5 | 20.2831 | TEa
Olanzapine | Psychotherapeutic drugs | Based on EMA/FDA | 50 | ug/L | 26.5 | 26.5779 | 13.25 | 13.2889 | SDsa1
Paroxetine | Psychotherapeutic drugs | Based on EMA/FDA | 100 | ug/L | 26.5 | 21.2392 | 26.5 | 21.2392 | TEa
Quetiapine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 16.6293 | 53 | 33.2586 | TEa
Risperidon | Psychotherapeutic drugs | Based on EMA/FDA | 20 | ug/L | 26.5 | 21.5917 | 5.3 | 4.3183 | TEa
Sertraline | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 18.8494 | 53 | 37.6988 | TEa
Venlafaxine | Psychotherapeutic drugs | Based on EMA/FDA | 200 | ug/L | 26.5 | 17.2272 | 53 | 34.4545 | TEa
Z-10-OH-Nortriptyline | Psychotherapeutic drugs | Based on EMA/FDA | 51 | ug/L | 26.5 | 23.1179 | 13.515 | 11.7901 | TEa
PT-INR | PT-INR for oral anticoagulation | Biological variation | 1 | {INR} | 5.272 | 12.6375 | 0.05272 | 0.12637 | SDsa1
Adenovirus antibodies | Respiratory serology |  |  | {titer}
Adenovirus IgA antibodies | Respiratory serology |  |  | [arb'U]/mL
Adenovirus IgG antibodies | Respiratory serology |  |  | [arb'U]/mL
Adenovirus IgM antibodies | Respiratory serology |  |  | IZS **
Chlamydia pneumoniae/psittacii antistoffen | Respiratory serology |  |  | {titer}
Chlamydia pneumoniae/psittacii IgA antistoffen | Respiratory serology |  |  | {titer}
Chlamydia pneumoniae/psittacii IgG antistoffen | Respiratory serology |  |  | {titer}
Chlamydia pneumoniae/psittacii IgM antistoffen | Respiratory serology |  |  | {titer}
Coxiella burnetii antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii antibodies fase 1 | Respiratory serology |  |  | {titer}
Coxiella burnetii antibodies fase 2 | Respiratory serology |  |  | {titer}
Coxiella burnetii fase 1 IgG antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii fase 1 IgM antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii fase 2 IgG antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii fase 2 IgM antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii IgA antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii IgG antibodies | Respiratory serology |  |  | {titer}
Coxiella burnetii IgM antibodies | Respiratory serology |  |  | {titer}
Influenza A antibodies | Respiratory serology |  |  | {titer}
Influenza A IgA antibodies | Respiratory serology |  |  | {ratio}
Influenza A IgG antibodies | Respiratory serology |  |  | {ratio}
Influenza A IgM antibodies | Respiratory serology |  |  | {titer}
Influenza B antibodies | Respiratory serology |  |  | {titer}
Influenza B IgA antibodies | Respiratory serology |  |  | {ratio}
Influenza B IgG antibodies | Respiratory serology |  |  | {ratio}
Influenza B IgM antibodies | Respiratory serology |  |  | {titer}
Legionella IgA antibodies | Respiratory serology |  |  | IZS **
Legionella IgG antibodies | Respiratory serology |  |  | IZS **
Legionella IgM antibodies | Respiratory serology |  |  | IZS **
Legionella IgTotaal antibodies | Respiratory serology |  |  | IZS **
Mycoplasma pneumoniae antibodies | Respiratory serology |  |  | {titer}
Mycoplasma pneumoniae IgA antibodies | Respiratory serology |  |  | {ratio}
Mycoplasma pneumoniae IgG antibodies | Respiratory serology |  |  | {ratio}
Mycoplasma pneumoniae IgM antibodies | Respiratory serology |  |  | {ratio}
Mycoplasma pneumoniae IgM antibodies (2e test) | Respiratory serology |  |  | {ratio}
Parainfluenza 1 antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 1 IgA antibodies | Respiratory serology |  |  | IZS **
Parainfluenza 1 IgG antibodies | Respiratory serology |  |  | {ratio}
Parainfluenza 1 IgM antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 1-3 antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 1-3 IgA antibodies | Respiratory serology |  |  | {ratio}
Parainfluenza 1-3 IgG antibodies | Respiratory serology |  |  | IZS **
Parainfluenza 1-3 IgM antibodies | Respiratory serology |  |  | IZS **
Parainfluenza 2 antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 2 IgA antibodies | Respiratory serology |  |  | IZS **
Parainfluenza 2 IgG antibodies | Respiratory serology |  |  | {ratio}
Parainfluenza 2 IgM antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 3 antibodies | Respiratory serology |  |  | {titer}
Parainfluenza 3 IgA antibodies | Respiratory serology |  |  | IZS **
Parainfluenza 3 IgG antibodies | Respiratory serology |  |  | {ratio}
Parainfluenza 3 IgM antibodies | Respiratory serology |  |  | {titer}
Past bij recente adenovirus infectie | Respiratory serology |  |  | Qualit.
Past bij recente Chlamydia infectie | Respiratory serology |  |  | Qualit.
Past bij recente influenza A infectie | Respiratory serology |  |  | Qualit.
Past bij recente influenza B infectie | Respiratory serology |  |  | Qualit.
Past bij recente Legionella infectie | Respiratory serology |  |  | Qualit.
Past bij recente Mycoplasma infectie | Respiratory serology |  |  | Qualit.
Past bij recente parainfluenza infectie | Respiratory serology |  |  | Qualit.
Past bij recente Q-koorts | Respiratory serology |  |  | Qualit.
Past bij recente RSV infectie | Respiratory serology |  |  | Qualit.
RSV antibodies | Respiratory serology |  |  | {titer}
RSV IgA antibodies | Respiratory serology |  |  | {ratio}
RSV IgG antibodies | Respiratory serology |  |  | {ratio}
RSV IgM antibodies | Respiratory serology |  |  | IZS **
Antibodies against CCP | Rheuma |  | 100 | [arb'U]/mL |  | 70.1703 |  | 70.1703 | SDsa1
antibodies against CCP ULN categorie | Rheuma |  |  | Qualit.
Reumafactor ULN categorie | Rheuma |  |  | Qualit.
Rheumatoid factor (IgM) | Rheuma |  | 200 | [IU]/mL |  | 68.7206 |  | 137.4411 | SDsa1
Acrosomale regio <= 2 vacuoles | Semen analysis |  |  | Qualit.
Acrosomale vacuoles < 20% kop-opp. | Semen analysis |  |  | Qualit.
Acrosoom 40-70% van kop-opp. | Semen analysis |  |  | Qualit.
Bevat geen knik | Semen analysis |  |  | Qualit.
Concentration | Semen analysis |  | 50 | 10*6/mL |  | 48.6051 |  | 24.3025 | SDsa1
Conclusion | Semen analysis |  |  | Qualit.
Dikte vrijwel gelijk over de hele lengte | Semen analysis |  |  | Qualit.
ERC < 1/3 van kop-opp. | Semen analysis |  |  | Qualit.
Geen post-acrosomale vacuoles | Semen analysis |  |  | Qualit.
Iets dunner dan het middenstuk | Semen analysis |  |  | Qualit.
Immotile | Semen analysis |  | 25 | % |  | 70.6732 |  | 17.6683 | SDsa1
Lengte as middenstuk = lengte as kop | Semen analysis |  |  | Qualit.
Lengte is circa 10x normale koplengte | Semen analysis |  |  | Qualit.
Morfologie (% normaal, WHO 2010) | Semen analysis |  | 5 | % |  |  |  |  | SDsa1
Non-progressive | Semen analysis |  | 25 | % |  | 130.4707 |  | 32.6177 | SDsa1
Ongeveer even lang als de kop | Semen analysis |  |  | Qualit.
Ovaal van vorm | Semen analysis |  |  | Qualit.
Progressive | Semen analysis |  | 25 | % |  | 76.2004 |  | 19.0501 | SDsa1
Rapid progressive motile (%A) | Semen analysis |  | 50 | % |  |  |  |  | SDsa1
Regelmatig van vorm | Semen analysis |  |  | Qualit.
Regelmatige, egale omtrek | Semen analysis |  |  | Qualit.
Total count | Semen analysis |  | 100 | [5 empty] | TEa
Soluble IL2 receptor | Soluble IL2 receptor |  | 1000 | pg/mL |  | 51.4849 |  | 514.8491 | SDsa1
Chlamydia trachomatis (real-time PCR) | STD External Molecular Quality Assessment Scheme |  | 20 | {Cq} |  |  |  |  | SDsa1
Neisseria gonorrhoeae (real-time PCR) | STD External Molecular Quality Assessment Scheme |  | 20 | {Cq} |  |  |  |  | SDsa1
Trichomonas vaginalis (real-time PCR) | STD External Molecular Quality Assessment Scheme |  | 20 | {Cq} |  | 26.9181 |  | 5.3836 | SDsa1
antibodies against Intrinsic factor | Stomach / Liver antibodies |  | 30 | k[arb'U]/L |  | 178.7271 |  | 53.6181 | SDsa1
antibodies against mitochondria | Stomach / Liver antibodies |  | 1600 | {titer} |  | 417.6819 |  | 6682.911 | SDsa1
antibodies against parietal cell | Stomach / Liver antibodies |  | 400 | {titer} |  | 369.7153 |  | 1478.8613 | SDsa1
antibodies against smooth muscle | Stomach / Liver antibodies |  | 300 | {titer} |  | 375.1351 |  | 1125.4053 | SDsa1
6-MMP | Thiopurines | Based on EMA/FDA | 30 | umol/L | 26.5 | 31.1418 | 7.95 | 9.3425 | SDsa1
6-TGN | Thiopurines | Based on EMA/FDA | 3.5 | umol/L | 26.5 | 46.3649 | 0.9275 | 1.6228 | SDsa1
Factor II 20210 | Thrombophilia (genetics) |  |  | Qualit.
Factor V Leiden | Thrombophilia (genetics) |  |  | Qualit.
antibodies against Tg | Thyroid Antibodies |  | 1000 | [arb'U]/mL |  | 81.91 |  | 819.1001 | SDsa1
antibodies against TPO | Thyroid Antibodies |  | 500 | [arb'U]/mL |  | 47.883 |  | 239.4149 | SDsa1
antibodies against TSH-R | Thyroid Antibodies |  | 15 | [arb'U]/L |  | 94.3028 |  | 14.1454 | SDsa1
Toxic substance | Toxicology | Based on EMA/FDA |  | Qualit.
Tryptase | Tryptase |  | 30 | ug/L |  |  |  |  | TE
3-Metoxytyramine | Tumor markers |  | 100 | pmol/L |  | 65.9738 |  | 65.9738 | SDsa1
AFP | Tumor markers | Biological variation | 30 | k[IU]/L | 17.643 | 14.0936 | 5.2929 | 4.2281 | TEa | 2023-02-23
CA-125 | Tumor markers | Biological variation | 50 | k[arb'U]/L | 13.913 | 14.6042 | 6.9565 | 7.3021 | SDsa1 | 2023-02-23
CA-15.3 | Tumor markers | Biological variation | 30 | k[arb'U]/L | 20.831 | 16.2971 | 6.2493 | 4.8891 | TEa
CA-19.9 | Tumor markers | Biological variation | 100 | k[arb'U]/L | 17.938 | 17.7014 | 17.938 | 17.7014 | TEa | 2023-02-23
Calcitonin | Tumor markers |  | 2.5 | pmol/L | 27.493 |  | 0.68732 |  | SDsa1 | 2023-02-23
CEA | Tumor markers | Biological variation | 10 | ug/L | 20.532 | 13.0369 | 2.0532 | 1.3037 | TEa | 2023-02-23
Chromogranin A | Tumor markers | Biological variation | 100 | ug/L | 20.73 |  | 20.73 |  | SDsa1 | 2022-01-01
Gastrin | Tumor markers |  | 100 | ng/L |  | 44.2087 |  | 44.2087 | SDsa1
HCG plus HCG-beta | Tumor markers |  | 100 | [IU]/L |  | 10.859 |  | 10.859 | SDsa1
HE4 | Tumor markers | Biological variation | 70 | pmol/L | 10.672 |  | 7.4704 |  | SDsa1 | 2023-02-23
Metanefrine | Tumor markers |  | 450 | pmol/L |  | 24.5414 |  | 110.4362 | SDsa1
Normetanefrine | Tumor markers |  | 600 | pmol/L |  | 25.7127 |  | 154.2762 | SDsa1
PSA total | Tumor markers | Biological variation | 6 | ug/L | 16.247 | 15.8885 | 0.97482 | 0.95331 | TEa | 2021-01-01
PSA, free | Tumor markers | Biological variation | 0.5 | ug/L | 17.543 | 18.182 | 0.087715 | 0.09091 | TEa | 2021-01-01
PSA, ratio free/total | Tumor markers |  | 0.1 | {ratio} |  | 30.2006 |  | 0.030201 | SDsa1
Serotonin | Tumor markers |  | 1000 | nmol/L |  |  |  |  | SDsa1
Tg-Autoantibodies | Tumor markers | Biological variation | 0.04 | k[arb'U]/L | 27.622 | 6901.8392 | 0.011049 | 2.7607 | SDsa1
Thyreoglobulin | Tumor markers | Biological variation | 16.6667 | ug/L | 28.226 | 23.1151 | 4.9745 | 3.8525 | Te | 2024-01-25
Urine element | Urine Sediment (qualtative) |  |  | Qualit.
Beta-carotene | Vitamin A/E and beta-carotene | Biological variation | 2 | umol/L | 43.098 | 29.7692 | 0.86196 | 0.59538 | TEa
Vitamin A | Vitamin A/E and beta-carotene | Biological variation | 2.5 | umol/L | 10.5 | 25.6083 | 0.2625 | 0.64021 | SDsa1 | 2022-01-01
Vitamin E | Vitamin A/E and beta-carotene | Biological variation | 20 | umol/L | 10.533 | 23.2437 | 2.1066 | 4.6487 | SDsa1 | 2022-01-01
Pyridoxalphosphate (PLP) | Vitamin B1/B6 | Biological variation | 100 | nmol/L | 26.362 | 25.2225 | 26.362 | 25.2225 | TEa
TPP (Th. Diphosphate) | Vitamin B1/B6 | Biological variation | 100 | nmol/L | 13.84 | 23.4223 | 13.84 | 23.4223 | SDsa1
1 Adjusted for the bias for SDsa versus Tea, see MUSE manual. | Muse manual
Φ incl. uncertainty of target value (Uref)
** {ratio}; {OD_unit}; [IU]/L; [IU]/mL; [arb'U]/mL; [arb'U]/L; {titer}; %{inhibition}; {other}; N/A
*** For the Vit D expert values, all SKML samples of 2024 are measured in one run by four different laboratories using a LC-MS/MS method. The median of the four results is used to define the expert values, see article below: Various calibration procedures result in optimal standardization of routinely used 25(OH)D ID-LC-MS/MS methods. Niek F. Dirks, Hubert W. Vesper, Antonius E. van Herwaarden, Jody M.W. van den Ouweland, Ido P. Kema, Johannes G. Krabbe, Annemieke C. Heijboer, Clinica Chimica Acta 462 (2016) 49–54.
**** For IGF-1 and GH expert values all SKML samples of 2024 in one run including harmonisation sample measured in duplo by four different laboratoria (iSYS, Immulite, Liaison en Cobas apparatuur). The median of the four resuts is used to define the expert values.
***** For the FT4 expert values all SKML samples of 2024 are measured in three different runs using the FT4 reference method from Radboud UMC. The mean is used to define the expert values. This FT4 reference method is based on Houcke et al. (2011) IFCC international conventional reference procedure for the measurement of free thyroxine in serum. Clin Chem Lab Med 49; 1275-1281. Unfortunately, the current measurement methods have not yet been standardized to the FT4 reference method. Therefore, the expert values are listed on the report but source target will be consensus - method group.
****** For the cortisol expert values all SKML samples of 2024 are measured in duplo in four different runs using UPLC-MSMS. Analyses is peformed by Erasmus MC.
******* For the testosterone expert values, all SKML samples from 2024 are measured in duplicate during 3 separate runs at Erasmus MC by UPLC-MSMS. The Erasmus MC testosterone assay is certified within the Hormone Standardization program of the CDC (https://www.cdc.gov/labstandards/pdf/hs/CDC_Certified_Testosterone_Assays-508.pdf) which guarantees accuracy and imprecision. At the same time a CDC reference panel is measured per run for verification.
Note on K2176: SA nu groter dan Tea was eerst andersom.

## Sample aspects
ANALYTE | SCHEME | SCORED ITEMS | TARGET VALUE TYPE | METROLOGICAL TRACEABILITY TARGET VALUE | COMMUTABILITY | HOMOGENITY | STABILITY | MATERIAL TYPE * | MATERIAL EXTRA INFO
Acanthamoeba PCR | Acanthamoeba Extern. Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation | Amoebae and purified DNA in stabilisation buffer. | Store at room temperature. Analyse samples soon after receipt.
ACE | ACE | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Acetone | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | calf serum. | The samples must be stored at -24 to -16°C upon receiving.
Ethanol | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Ethylene glycol | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
GHB | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Isopropanol | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Methanol | Alcohols/GHB | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Type III allergie IgG AB (Qual.) | Allergy Type III antibodies | na | Consensus |  | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Frozen Human pooled serum diluted in omniplasma | The samples must be stored at -24 to -16°C upon receiving. (samples are shipped without specified shipping conditions)
Type III allergy IgG AB | Allergy Type III antibodies | Quant. | Consensus | Consensus - method group | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
ASXL1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Human bone marrow (gDNA & cDNA). | The ampoules arrive at room temperature. Genomic DNA should be stored at 4ºC and cDNA at -20ºC upon arrival
BCOR mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CBFB-MYH11 | AML molecular diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CEBPA bZIP mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CEBPA dm | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
EZH2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
FLT3 ITD | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
FLT3 TKD | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
GATA2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IDH1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IDH2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
KIT mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
NPM1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
RUNX1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
RUNX1-RUNX1T1 | AML molecular diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
SF3B1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
SRSF2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
STAG2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
TP53 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
U2AF1 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
ZRSR2 mutation | AML molecular diagnostics | Quant. and Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Acetaminophen | Analgesic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Ibuprofen | Analgesic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Salicylic acid | Analgesic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
ANA IIF interfererend met ANCA | ANCA-GBM | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Human plasmapheresis material or serum. | The liquid samples should be stored in the freezer -24°C to -16°C.
ANCA IIF (ethanol gefixeerd) | ANCA-GBM | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
ANCA IIF (formaline ANCA IIF (formaline gefixeerd) | ANCA-GBM | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
ANCA IIF conclusie | ANCA-GBM | Qualit. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against GBM | ANCA-GBM | Qualit. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against MPO | ANCA-GBM | Qualit. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against PR3 | ANCA-GBM | Qualit. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Sterkte van het patroon | ANCA-GBM | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Amikacin | Antibiotic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Gentamicin | Antibiotic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Methotrexate | Antibiotic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Tobramycin | Antibiotic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Vancomycin | Antibiotic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Ceftazidim | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Frozen Human EDTA plasma. | The samples must be stored at <-70°C upon receiving
Ciprofloxacine | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Flucloxacilline | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
N-acetyl Sulfamethoxazol | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Piperacilline | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
protein-unbound (free) flucloxacillin | Antibiotic drugs 2 | Quant. | Consensus | Consensus - method group | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sulfamethoxazol | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Trimethoprim | Antibiotic drugs 2 | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically verified with accepted imperfections | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Antibodies against desmogleïne 1 | Antibodies against skin | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Upon receiving, store samples at -20 to -80°C until time of analysis.
Antibodies against desmogleïne 3 | Antibodies against skin | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Antibodies against NC16A domain BP180 | Antibodies against skin | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Basement membrane (BMZ) | Antibodies against skin | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
extracellu. surface keratinocytes (ECS) | Antibodies against skin | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Carbamazepine | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Epoxide-cbz | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Ethosuximide | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Lamotrigine | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Levetiracetam | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Mono-hydroxy-carbazepine | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Phenobarbital | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Phenytoin | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Valproic acid | Antiepileptic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Phenytoin | Antiepileptic drugs free concentration | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Phenytoïn free concentration | Antiepileptic drugs free concentration | na | Consensus | Consensus - method group | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Valproic acid | Antiepileptic drugs free concentration | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Valproic acid free concentration | Antiepileptic drugs free concentration | na | Consensus | Consensus - method group | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Fluconazole | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Flucytosine | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxy-itraconazole | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Isavuconazol | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Itraconazole | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Posaconazole | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Voriconazole | Antifungal drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Amprenavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Atazanavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Darunavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Dolutegravir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Efavirenz | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Elvitegravir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Etravirine | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Lopinavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Nevirapine | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Raltegravir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Rilpivirine | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Saquinavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Tipranavir | Antiretroviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
acetyl-isoniazide | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Frozen Human Li-heparinized plasma and millipore water. | The samples must be stored at <-70°C upon receiving.
Bedaquiline | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
clofazimine | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Delamanid | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desacetylrifabutin | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Ethambutol | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Isoniazid | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Levofloxacin | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Linezolid | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Moxifloxacin | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Pyrazinamide | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
rifabutin | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Rifampicin | Antituberculosis drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Aciclovir | Antiviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised calf serum. | The samples must be stored at 2 - 8°C upon receiving.
Ganciclovir | Antiviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Ribavirin | Antiviral drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
AF-bone | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
AF-Gal | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
AF-intestinal | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
AF-liver 1 | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
AF-liver 2 | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
AF-Macromoleculaire fraction | AP-isoenzymes | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Other | AP-isoenzymes | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Anti-DNaseB | AST / AntiDNaseB serology | na | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
AST | AST / AntiDNaseB serology | na | Expert | Analytical by 2 experts | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Other(specify at comments) | AST / AntiDNaseB serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Wijst op recente streptokokkeninf. | AST / AntiDNaseB serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Determination Bacteriology | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. clindamycine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. clindamycine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. meropenem (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. metronidazole (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. metronidazole (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. piperacilline/tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Bac. piperacilline/tazobactam (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Camp. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Cory. Rifampicin (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Cory. Rifampicin (report) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Kleb. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Kluy. Piperacillin/Tazobactam (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Kluy. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Sta. aur. Mupirocine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Sta. aur. Mupirocine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Sta. spp. oxacilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Sta. spp. oxacilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. clindamycine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. clindamycine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Co-trimoxazol (Raporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. doxycycline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. erythromycine (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. erythromycine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Penicilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Penicilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. tetracycline (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Vancomycin (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
G. Str. Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Aci. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Cit. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Cit. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Cit. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E. faecium High level Gentamicin | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E. faecium resistance (HLGR) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E. faecium VRE | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E.faecalis High level Gentamicin | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E.faecalis resistance (HLGR) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. E.faecalis VRE | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Entbac. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Entbac. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Entbac. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Esch. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Esch. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Esch. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Haf. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Haf. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Haf. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kleb. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kleb. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kleb. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kluy. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kluy. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Kluy. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Lecl. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Lecl. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Lecl. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Morg. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Morg. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Morg. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prot. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prot. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prot. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prov. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prov. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Prov. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Pseu. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Salm. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Salm. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Salm. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Ser. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Ser. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Ser. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Shig. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Shig. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Shig. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Sta. aur. MLSB inducibale | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Sta. aur. MRSA | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Sta spp. MLSB inducible | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Str. MLSB inducible | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Str. pne. MLSB inducible | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Yer. Carbapenemase | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Yer. ESBL | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
RM. Yer. Plasmid AmpC | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Amikacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Amikacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Colistin (Reportin) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Levofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Levofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aci. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Ampicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Ampicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Nitrofurantoin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Nitrofurantoin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Norfloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Norfloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Penicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aec. Penicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Aztreonam (measuring) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Aztreonam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Cefepime (measuring) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Levofloxacine (measuring) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Aer. Levofloxacine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Bac. amoxicilline/clavulaanzuur (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Bac. amoxicilline/clavulaanzuur (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Azithromycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Azithromycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Clarithromycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Clarithromycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Doxycycline (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Doxycycline (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Erythromycin (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Erythromycin (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Camp. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cit. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Clos. Metronidazole (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Clos. Metronidazole (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Clos. Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Clos. Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Clindamycin (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Clindamycin (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Penicillin (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Penicilline (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Cory. Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Ampicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Ampicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Nitrofurantion (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Nitrofurantion (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Teicoplanin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Teicoplanin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecalis Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Ampicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Ampicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Teicoplanin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Teicoplanin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. E.faecium Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Entbac. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Cefuroxime (Rapportage) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Nitrofurantoin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Nitrofurantoin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Esch. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ampicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ampicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Rifampicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Rifampicin (Rapportage) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. inf. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ampicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ampicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Hae. par. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Meropenem (Rapportage) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Haf. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Augmentin (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Cefepime (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Cefepime (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Cefotaxime (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kleb. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Kluy. Tobramycin (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Cefuroxime (Raporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lecl. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. amoxicilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. amoxicilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. ampicilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. ampicilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. erythromycine (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. erythromycine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. meropenem (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. penicilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. penicilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. trimethoprim-sulfamethoxazole (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Lis. trimethoprim-sulfamethoxazole (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Morg. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Azithromycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Azithromycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. gon. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Penicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Penicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Rifampicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Neis. men. Rifampicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Penicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Penicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Past. mul. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Augmentin (Repporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefotaxim (Reporting) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefotaxime (measured) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Co-trimoxazol (Raporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prot. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Cefuroxime (Raporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Piperacilline/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Prov. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Ciprofloxacine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Colistin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Colistin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Piperacillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Piperacillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Pseu. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefepime (Rapportage) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefuroxime (measuerd) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Salm. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Ser. Tobramycine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Shig. Tobramycine (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Clindamycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Clindamycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Erythromycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Erythromycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Flucloxacillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Flucloxacillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Rifampicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Rifampicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. aur. Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. pseudi. oxacilline (Measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. pseudi. Oxacilline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Clindamycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Clindamycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Erythromycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Erythromycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Flucloxacillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Flucloxacillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Rifampicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Rifampicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Tetracyclin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Tetracycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Vancomycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Sta. spp. Vancomycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Doxycycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Doxycycline (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Penicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Str. pne. Penicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Cefotaxime (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Ceftazidime (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Chlooramphenicol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Chlooramphenicol (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Ciprofloxacine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Ciprofloxacine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Colistine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Colistine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Co-trimoxazol (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Erytromycine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Erytromycine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Meropenem (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Neomycine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Neomycine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Piperacilline/ Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Piperacilline/ Tazobactam (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Tetracycline (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Tetracycline (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Trobramycine (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Vib. Trobramycine (reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Amoxicillin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Amoxicillin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Augmentin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Augmentin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefepime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefepime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefotaxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefotaxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ceftazidime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ceftazidime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ceftriaxone (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ceftriaxone (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefuroxime (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Cefuroxime (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ciprofloxacin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Ciprofloxacin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Co-trimoxazol (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Co-trimoxazol (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Gentamicin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Gentamicin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Imipenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Imipenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Meropenem (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Meropenem (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Piperacillin/Tazobactam (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Piperacillin/Tazobactam (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Tobramycin (measured) | Bacteriology | Qualit. | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
S. Yer. Tobramycin (Reporting) | Bacteriology | na | Expert | Analytical by 5 experts | Analytically different | Analytically different | Analytically different | Freeze dried material. | The samples must be stored at 2 - 8°C upon receiving.
B. henselae IgG | Bartonella henselae serology | Quant | [5 empty] | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
B. henselae IgM | Bartonella henselae serology | Quant
B. henselae Igtotaal | Bartonella henselae serology | Quant
Matches recent Bartonella henselae infection | Bartonella henselae serology | Qualit.
Clobazam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised calf serum. | The samples must be stored at 2 - 8°C upon receiving.
Clonazepam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylclobazam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethyldiazepam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Diazepam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxymidazolam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Midazolam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Oxazepam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Temazepam | Benzodiazepines | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Total Bile Acid Salts | Bile acid salts | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Human Lyophilised serum and Li-heparanized plasma. | The samples must be stored at 2 - 8°C upon receiving.
Adalimumab | Biologicals | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Adalimumab antibodies | Biologicals | na | Expert | Analytical by expert (RIA) | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Infliximab | Biologicals | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Infliximab antibodies | Biologicals | na | Expert | Analytical by expert (RIA) | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Vedolizumab | Biologicals | na | Expert | Analytical by expert (RIA) | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Abn.lymphocyte susp.malignant | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Abn.lymphocyte susp.reactive | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Agglutination | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Alder-Reilly anomaly | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Atypical lymphocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Auer rods | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Basophilic granulocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Blasts | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
cleaved lymfocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Döhle bodies | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Echinocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Elliptocytes | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Eosinophilic granulocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Erythroblasts (outside WBC count) | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Giant granulocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Gumprecht shadows (VHL) | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Hairy cells | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Howell Jolly bodies | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Hypersegmentation | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Hypochromasia | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Hypogranulation | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Hypogranulation | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
LGL cells > 20% of >2.0x10 | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
lymfocytes grumelee core pattern | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Lymphocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Macrocytosis | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Macropolycyte | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Macrothrombocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Malaria | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Maturation dissociation | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Megakaryocytes remnants | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Megalocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Metamyelocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Microcytosis | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Monocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Myelocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Neutroph. granulocytes (bands) | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Neutroph. granulocytes (PMNS) | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Not to differentiate | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Plasma cells | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Plasmablasts | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Poikylocytose | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Polychromasia | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Prolymfocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Promonocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Promyelocytes | Blood Cell Morphology | % | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Pseudo Pelger Huet | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Punct.Basophilia | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Rouleaux formation | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Satellitism | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Schizocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Sickle cells | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
SLVL cells | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Spherocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Stomatocytes | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Target cells | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Tear drop cells | Blood Cell Morphology | Quant. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Thrombocyte aggregates | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Toxic granulation | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
Vacuolisation | Blood Cell Morphology | Qualit. | Expert | Analytical by 4 experts | unknown | Analytically different | Analytically different | Periferal human bloodsmear. | Store at room temperature upon receiving.
P.species antigen rapid test | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different | Human EDTA blood and blood lysate. | The samples must be stored at 4°C upon receiving.
Conclusion malaria antigentest | Blood parasites | Qualit. | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
Malaria rapid test | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
Moleculaire sneltest (LAMP) | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
Molecular research Malaria | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. falciparum PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. knowlesi PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. malariae PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. ovale PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. species PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P. vivax PCR | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
P.falciparum antigen rapid test | Blood parasites | na | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different
Parasite in blood | Blood parasites | Qualit. | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different | Giemsa coloured human blood smear. | Store at room temperature upon receiving.
Pf/Pk parasitemie | Blood parasites | Quant. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Bilirubin | Blood pigments in cerebrospinal fluid | na | Expert | Edited consensus by expert (median) | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen diluted human serum. | The samples must be stored at -24 to -16°C upon receiving. (samples are shipped without specified shipping conditions)
Bilirubin (absorption) | Blood pigments in cerebrospinal fluid | na | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Most appropriate conclusion with regard to SAB | Blood pigments in cerebrospinal fluid | Qualit. | Expert | Analytical by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Oxyhemoglobin | Blood pigments in cerebrospinal fluid | na | Expert | Edited consensus by expert (median) | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Oxyhemoglobin (absorption) | Blood pigments in cerebrospinal fluid | na | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Ca++ (actual) | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Aqueous solution with bovine albumin. | The samples must be stored in dark, at 2 - 8°C upon receiving.
Cl- | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Glucose | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
K+ | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Lactate | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Na+ | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
PCO2 | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
pH | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
PO2 | Bloodgas Parameters | na | Consensus | Consensus - method group | probably not commutable | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
(clusters of) tumor cells | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Aantal megakaryocyten pgz (100x) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Alternative bone marrow diagnosis 1 | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Alternative bone marrow diagnosis 1 | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Amyloïd | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Auer rods | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Band form (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Basophil (bone marrow) | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Blast incl. promonocyten (bonemarrow) | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Bone marrow particles | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Cellularity | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Cytogenetics | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Cytoplasm. Vacuolisatie | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Dysplasie quantitative (Erytropoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Dysplasie quantitative (Granulopoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Dysplasie quantitative (Megakaryopoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Döhle bodies | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Eosinophil (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Erytroblast (bonemarrow) | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Haard(en) >50 cellen | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Hemofagocytose | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Histologiy (biopsy) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Hyperlobulaire vormen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Hypersegmentation | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Hypogranulation | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Hypolobulaire vormen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Immunophenotyping | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Iron incorporation (sideroblasts) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Iron stock | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Karyorrhexis | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kernbruggen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kernmisvormingen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kwantiteit (Erytropoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kwantiteit (Granulopoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kwantiteit (Lymfopoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Kwantiteit (Plasmacellen) | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Losliggende kernen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Lymphocyte (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
M-E ratio | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Macrofagen toegenomen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Meerkernig | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Meest waarschijnlijke beenmerg diagnose | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Megaloblastair | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Metamyelocyte (bone marrow) | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Microbiological examination | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Micromegakaryocyten | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Molecular research | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Monocyte (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Monopoiesis (morfology) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Morfologie (Plasmacellen) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Myelocyte (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Neutrophil (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
None | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Nuclear malformations | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Ongebruikelijke kleine vormen | Bonemarrow morphology | Quant. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Ongebruikelijke reuze vormen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Other | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Pathologische mestcellen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Percentage promonocyten | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Plasma cell (bone marrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Promyelocyt (bonemarrow) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Pseudo Chédiak Higashi granules | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Pseudo Pelger-HUet granules | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Pseudo-Gaucher cellen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Representative | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Ringsideroblasten % | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Sea blue histiocytes | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Suspect malignant | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Suspect reactive | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Toename mestcellen | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Uitrijping (Erytropoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Uitrijping (Granulopoëse) | Bonemarrow morphology | Qualit. | Expert | Analytical by 6 experts | NA | NA | NA | Digital image. | NA
Bordetella IgG | Bordetella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Past bij recente Kinkhoest infectie | Bordetella serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella Agglutination | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
Brucella IgG (qual.) | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella IgG [arb'U]/mL | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella IgG {ratio} | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella IgM (qual.) | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella IgM [arb'U]/mL | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Brucella IgM {ratio} | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Other(specify at comments) | Brucella serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij infectie met Brucella | Brucella serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Busulfan | Busulfan | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Retrospectively based on between-lab variation | Analytically different | Busulfan in 0.1 ml organic solvent, add 1.00 ml calfserum (included). | The samples must be stored at -20°C or below upon receiving. (samples will be send cooled 2-10 °C)
Amiodarone | Cardiac drugs | Quant. | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised calf serum. | The samples must be stored at -16 to -20°C upon receiving.
Caffeine | Cardiac drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desethylamiodarone | Cardiac drugs | Quant. | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Digoxin | Cardiac drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Flecainide | Cardiac drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Theophylline | Cardiac drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
BNP | Cardiac markers | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | The samples must be stored by < -70°C upon receiving. (samples are shipped on dry-ice)
CKMB act | Cardiac markers | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CKMB mass | Cardiac markers | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
hs-Troponin I | Cardiac markers | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
hs-Troponin T | Cardiac markers | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Myoglobin | Cardiac markers | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
NT-proBNP | Cardiac markers | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CDT | CDT | quant. | Reference | Reference method according JCTML list | Analytically different | Retrospectively based on between-lab variation | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | The samples must be stored by < -70°C upon receiving. (samples are shipped on dry-ice)
CDT, qualitative | CDT | Qualit. | Consensus | Consensus - method group | Analytically different | Retrospectively based on between-lab variation | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Anti-TG2 IgA ULN categorie | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | The samples must be stored at -20°C upon receiving. (samples are shipped without specified shipping conditions)
IgA against endomysium | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA against gedeamideerd gliadine | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA against tissue transglutaminase | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG against gedeamideerd gliadine | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG tegen endomysium | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG tegen tissue transglutaminase | Celiac disease | na | Consensus | Consensus - method group | probably not commutable | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Donor 1 cellen | Chimerism | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human K3EDTA blood and DNA. | Store samples at room temperature and DNA at 2 - 8°C upon receiving.
Donor 2 cellen | Chimerism | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Patiënt cellen | Chimerism | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
ALAT | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | The samples must be stored in the freezer (< -70°C) upon receiving. Once thawed, the samples are stable during 6 hours at 6 - 24°C. When stored at 2 - 8°C the samples are stable during 24 hours (in the light protective brown glass vial).
Albumin | Clinical Chemistry, blood | na | Expert | Value assignment by expert lab using reference-material (ERM-DA470k) as calibrator | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Alk. Phosphatase | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Ammonia | Clinical Chemistry, blood | Quant. | Expert | ALTM | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Amylase | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
ASAT | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Bilirubin | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Bilirubin direct | Clinical Chemistry, blood | Quant. | Expert | ALTM | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Calcium | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Chloride | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CK | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Corrected Calcium | Clinical Chemistry, blood | na | Consensus | Consensus - method group | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Creatinine | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
eGFR (F, 55, white) | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Gamma-GT | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Glucose | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Inorg. Phosphate | Clinical Chemistry, blood | Quant. | Expert | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Iron | Clinical Chemistry, blood | Quant. | Expert | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Lactate | Clinical Chemistry, blood | Quant. | Expert | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
LD | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Lipase | Clinical Chemistry, blood | Quant. | Expert | ALTM | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Lithium | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Magnesium | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Osmolality | Clinical Chemistry, blood | Quant. | Expert | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Potassium | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Pseudo Cholinesterase | Clinical Chemistry, blood | Quant. | Expert | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Sodium | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Protein | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urate | Clinical Chemistry, blood | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urea | Clinical Chemistry, blood | Quant. | Expert | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Acetylacetate | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Human urine. | The samples must be stored at 2 - 8°C upon receiving. Sample (2ml) in ACU-DROP-II dripper, measure as soon as possible.
Bilirubin | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Erytrocytes/Hb | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Glucose | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ketones | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Leucocytes | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Nitrite | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
pH | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Pregnancy | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Protein | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
S.G. | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Urobilinogen | Clinical Chemistry, urine qualitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
(micro)Albumin | Clinical Chemistry, urine quantitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human urine. | The samples must be stored at 2 - 8°C upon receiving.
Amylase | Clinical Chemistry, urine quantitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Calcium | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Chloride | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Creatinine | Clinical Chemistry, urine quantitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Glucose | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Inorg. Phosphate | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Magnesium | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Osmolality | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Oxalate | Clinical Chemistry, urine quantitative | na | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Potassium | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Sodium | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Protein | Clinical Chemistry, urine quantitative | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urate | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urea | Clinical Chemistry, urine quantitative | Quant. | Expert | Analytical by expert | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CMV IgG | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
CMV IgM | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CMV-IgG aviditeit | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
EBNA IgG | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
EBV-EA IgG | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
EBV-EA IgM | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
EBV-VCA IgG | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
EBV-VCA IgM | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Heterof a.s./slide tst/Paul&Bunell | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Other(specify at comments) | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij doorgemaakte CMV-inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij doorgemaakte EBV-inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij doorgemaakte Toxoplasma inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente CMV-inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente EBV-inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente Toxoplasma-inf. | CMV/EBV/Toxoplasma serology | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Toxoplasma IgG | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Toxoplasma IgG aviditeit | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Toxoplasma IgM | CMV/EBV/Toxoplasma serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Antitrombin | Coagulation | Quanti | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different | Lyophilised human citrate plasma. | The samples must be stored at 2 - 8°C upon receiving.
APTT | Coagulation | na | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
Factor VIII | Coagulation | Quanti | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
Fibrinogen | Coagulation | na | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
PT | Coagulation | na | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
PT-INR | Coagulation | na | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
ANA (HEp2 IIF screen) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | The samples must be stored at 2 - 8°C upon receiving.
antibodies against dsDNA | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Antibodies against ENA (solid phase screen) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CenpB | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Dominant HEp2 IIF pattern | Collagen | na | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Histonen | Collagen | na | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Jo-1 | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Number of HEp2 IIF patterns | Collagen | na | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Ribosomaal P protein (rRNP) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
RNP | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Sm | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
SS-A (52 + 60 kD) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
SS-A (52kD) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
SS-A (60 kD) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
SS-B | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Topoisomerase I (Scl70) | Collagen | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
AP | Complement factors | Qualit. | Expert | Theoretical by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Store at < -20°C upon receiving. (samples are shipped on dry-ice)
C1q | Complement factors | Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
C3 | Complement factors | Quant. and Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
C4 | Complement factors | Quant. and Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CP | Complement factors | Qualit. | Expert | Theoretical by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
MBL | Complement factors | Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
MBL-P | Complement factors | Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against GAD | Diabetes | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Store at -16 to -24°C upon receiving. (samples are shipped without specified shipping conditions)
antibodies against IA2 | Diabetes | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against ICA | Diabetes | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Acenocoumarol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Human urine, or calf serum. | See instructions
Adalimumab | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Aloe-emodin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Amoxicillin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Atenolol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Baclofen | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Benzylpenicillin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Bisacodyldiphenol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Bisoxantindiphenol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Bupropion | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Cefuroxime | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Chloortalidon | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Chlorothiazide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Chrysophanic acid | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Clopentixol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Danthron | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylflunitrazepam | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylmaprotiline | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Disopyramide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Duloxetine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Emodin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Felbamaat | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Fenprocoumon | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Flucloxacillin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Flunitrazepam | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Flupenthixol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Furosemide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Gabapentine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Glibenclamide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Haloperidol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydrochlororthiazide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxybupropion | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxychloroquine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxymidazolam | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxymidazolamglucuronide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Infliximab | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Kinidine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Lacosamide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Maprotiline | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
mefruside | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
meropenem | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Metformin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Metoprolol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Midazolam | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
moclobemide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Mycophenolic acid | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
N4-acetyl-sulfamethoxazole | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
N-acetyl-procainamide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Penfluridol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Phenolphtaleïne | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Pimozide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Pipamperon | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Pregabaline | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Procaïnamide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Propranolol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Rhein | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sertindol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sotalol | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Spironolacton | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sulfamethoxazole | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Teicoplanine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Tolbutamide | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Topiramate | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
tranylcypromine | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Triamterene | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Trimethoprim | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Warfarin | Drugs alternating | na | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Acetaminophen confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Human urine. | Store at -16 to -24°C upon receiving.
Acetaminophen screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Alfa-PVP screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Alfa-PVP confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Alprazolam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Amphetamine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Amphetamine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Aspect urine | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Barbiturate confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Barbiturate screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Benzodiazepine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Benzoylecgonine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Bromazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Buprenorphine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Buprenorphine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cannabis screeninng | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Clonazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cocaïne confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cocaïne screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Codeine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Creatinine | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Diazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Drugs | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
EDDP confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
EDDP screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ethanol confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ethanol screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ethylglucuroninde confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ethylglucuroninde screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Fentanyl screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Fentanyl confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Flunitrazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Flurazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
GHB screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
GHB confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ketamine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ketamine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Lorazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Lormetazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
LSD confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
LSD screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
MDA confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
MDEA confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
MDMA confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methadone confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methadone screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methylamphetamine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methylfenidaat screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methylfenidaat Confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Midazolam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Mitragynine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Mitragynine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Morphine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Nitrazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Norcodeïne confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Nordiazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Opiates screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Oxazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Phenazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Phencyclidine screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Phencyclidine confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Propoxyfeen screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Propoxyfeen confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
TCA confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
TCA screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Temazepam confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
THCCOOH confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Tramadol screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Tramadol confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
XTC/methamfetaminen screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
3-MMC screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
3-MMC confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
4-FA screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
4-FA confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
4-MMC/mefedron screening | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
4-MMC/mefedron confirmation | Drugs of Abuse | Qualit. | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
3 celspanel | Erythrocyte serology basic | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different | Human SAGM erytrocytes and CPD plasma. | Store at 4°C upon receiving.
ABO conclusie | Erythrocyte serology basic | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Anti-A | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Anti-B | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Anti-D | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Anti-D controle | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
A-testerytrocytes | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
B-testerytrocytes | Erythrocyte serology basic | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Kruis. donor 1 | Erythrocyte serology basic | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Kruis. donor 2 | Erythrocyte serology basic | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Rhesus-D conclusie | Erythrocyte serology basic | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Anti C | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different | Human serum. | Store at 2 - 8°C upon receiving.
Anti c | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti C(w) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti D | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti E | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti e | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti f | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Fy(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Fy(b) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Jk(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Jk(b) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti K | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti k | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Kp(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Le(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Le(b) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Lu(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti M | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti N | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti P1 | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti S | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti s | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
Anti Wr(a) | Erythrocyte serology detailed | Qualit. | Expert | Analytical by 3 experts (using standard erythrocyte serology methods) | Theoretically likely | Analytically different | Analytically different
3 celspanel | Erythrocyte serology pregnancy | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human SAGM erytrocytes and CPD plasma. Human CPD cord blood from a newborn | Store at 4°C upon receiving.
ABO conclusie | Erythrocyte serology pregnancy | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Anti-A | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Anti-B | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Anti-c | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Anti-D | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Anti-D controle | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
A-testerytrocytes | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
B-testerytrocytes | Erythrocyte serology pregnancy | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Rhesus-c conclusion | Erythrocyte serology pregnancy | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Rhesus-D conclusie | Erythrocyte serology pregnancy | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Ethanol | Ethanol in serum | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | Store at < -70°C (samples are shipped on dry-ice)
Doorgemaakte/vaccinatie Mazelen inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
Doorgemaakte/vaccinatie Rubella inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Herpes simplex antibodies | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Herpes simplex IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Herpes simplex IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Herpes simplex type 2 IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Herpes simplex type 2 IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
HHV-6 IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
HHV-6 IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Measels antibodies | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Measels IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Measels IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Other(specify at comments) | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Parvovirus B19 IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Parvovirus B19 IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij doorgemaakte HSV inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij doorgemaakte parvo B19 inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij doorgemaakte VZV inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij recente HSV inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij recente Mazelen inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij recente Parvo. B19 inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij recente Rubella inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Past bij recente VZV inf. | Exantheme Diseases Diagnostics | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Rubella IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Rubella IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Varicella zoster antibodies | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Varicella zoster IgG | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Varicella zoster IgM | Exantheme Diseases Diagnostics | na | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Alpha-1-antitrypsin | Feces-chemistry | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human feces. | Store at -16 to -24°C upon receiving. (samples are shipped on dry-ice)
Bile acids | Feces-chemistry | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Calprotectin | Feces-chemistry | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Elastase | Feces-chemistry | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Fat in feces | Feces-chemistry | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Fetal fibronectin | Fetal fibronectine | na | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Human vaginal fluid | Store at 2 - 8°C upon receiving. Measure as soon as possible.
Fetal cells | Fetomaternal transfusion | na | Consensus | ALTM | NA | Analytically different | Analytically different | Human K3/Na3EDTA blood | Store at 4°C upon receiving. Measure as soon as possible.
Foetomaternale hemorrhage | Fetomaternal transfusion | na | Consensus | ALTM | NA | Analytically different | Analytically different
Fructosamin | Fructosamin | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | Store at < -70°C upon receiving. (samples are shipped on dry-ice)
HbA | Hb variants | Quant. | Expert | Analytical by expert (HPLC methode) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen whole human blood. | Store at < -70°C upon receiving.(samples are shipped on dry-ice)
HbA2 | Hb variants | Quant. | Expert | Analytical by expert (HPLC methode) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HbF | Hb variants | Quant. | Expert | Analytical by expert (HPLC methode) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HbX1 | Hb variants | Quant. | Expert | Analytical by expert (HPLC methode) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HbX2 | Hb variants | Quant. | Expert | Analytical by expert (HPLC methode) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HbA1c | HbA1c whole blood | Quant. | Reference | Reference method according JCTML list | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Human whole blood. | Store at 2 - 8°C upon receiving. Measure as soon as possible.
Helicobacter pylori Antigeen in Feces | Helicobacter pylori (Ag) in feces | Qualit. | Consensus | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human feces. | Store at 2 - 8°C upon receiving.
Helicobacter antibodies | Helicobacter serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serm | Store at 2 - 8°C upon receiving.
Helicobacter IgA | Helicobacter serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Helicobacter IgG | Helicobacter serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Helicobacter serologie past bij | Helicobacter serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Other(specify at comments) | Helicobacter serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Ancylostoma spp PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation | Human feces (ethanol preserved) and purified DNA. | Store feces at room temperature at room temperature and DNA at 4°C upon receiving. Measure as soon as possible.
Ascaris spp PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Necator americanus PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Schistosoma spp PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Strongyloides stercoralis PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Trichuris trichiura PCR | Helminth External Mol. Quality Assessment | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Erytrocytes | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation | Fresh human blood. | Store at room temperature (5 - 25°C) upon receiving. The samples are stable only on the day of measurement.
Hematocrit | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Hemoglobin | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
IPF | Hemocytometry | Quant. | Consensus | ALTM | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Leucocytes | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Lymfocytes abs. | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
MCV | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Monocytes abs. | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Neutrofiles abs. | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
RDW-SD | Hemocytometry | Quant. | Consensus | ALTM | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
RET-He | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Reticulocytes abs. | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Trombocytes | Hemocytometry | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Retrospectively based on between-lab variation
Fraction COHb | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophillised Bovine blood. | Store at 2 - 8°C upon receiving.
Fraction HHb | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Fraction MetHb | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Fraction O2Hb | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
sO2 | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Hb | Hemoximetry | na | Consensus | Consensus - method group | probably not commutable | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Beschermd tegen hepatitis A? | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophillised human serum. | Store at 2 - 8°C upon receiving.
Beschermd tegen hepatitis B? | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis A IgG | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis A IgM | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis A IgT | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B Anti-HBcore | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B Anti-HBe | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B Anti-HBs | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B HBeAg | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B HBsAg | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B HBsAg confirmation | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B HBsAg kwantitatief | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis B IgM anti-HBcore | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti C1 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti c100p + 5-1-1p | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti C2 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti c22p | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti C33c | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti NS3 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti NS4 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C anti NS5 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C Anti-HCV | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C Anti-HCV confirmation | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis C-anti E2 | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis E IgG | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hepatitis E IgM | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Other(specify at comments) | Hepatitis ABCE serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
past bij acute infectie hep A | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij acute infectie hep E | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
past bij acute/chronisch inf. hep B | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
past bij doorgemaakte hep B inf. | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij infectie hep C | Hepatitis ABCE serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-gp105 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophillised human serum. | Store at 2 - 8°C upon receiving.
anti HIV-gp120 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-gp160 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-gp36 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-gp41 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p17 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p24 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p31 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p39 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p51 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p55 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
anti HIV-p66 | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Fits infection with HIV | HIV serology (incl. blot) | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HIV p24 antigentest | HIV serology (incl. blot) | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HIV screening (1e test) | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HIV screening (2e test) | HIV serology (incl. blot) | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HIV-confirmation (immuneblot) | HIV serology (incl. blot) | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij acute infectie met HIV | HIV serology (incl. blot) | Qualit. | Expert | Analytical by expert | NA | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Homocystein | Homocystein | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
11-Deoxycortisole | Hormones in Saliva | na | Consensus | Consensus - method group | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Frozen human saliva. | Store at < -70°C upon receiving. (samples are shipped on dry-ice)
21-deoxycortisole | Hormones in Saliva | na | Consensus | Consensus - method group | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Androsteendion | Hormones in Saliva | Quant. | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cortisol | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cortison | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
DHEA | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Melatonin | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Progesteron,17OH | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Testosterone | Hormones in Saliva | na | Consensus | ALTM | Analytically different | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
11 deoxycorticosterone | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different | Frozen human serum. | Store at < -70°C upon receiving. (samples are shipped on dry-ice)
11-Deoxycortisol | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
21-deoxycortisol | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
25 OH Vitamin D | Hormones in Serum | Quant. | Expert *** | Analytical by 4 experts | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
ACTH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Active B12 | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Aldosteron | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
AMH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Androsteendion | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Copeptin | Hormones in Serum | na | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Corticosterone | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Cortisol | Hormones in Serum | Quant. | Expert ****** | Analytical by expert | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
C-peptide | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
CTX | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
DHEA | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
DHEAS | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Digoxin | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Eestrone | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Ferritin | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Folate | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Free testosteron | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
FSH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Growth hormone | Hormones in Serum | Quant. | Expert **** | Analytical by 4 experts | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
IGF-1 | Hormones in Serum | Quant. | Expert **** | Analytical by 4 experts | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
IGF1-SD | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Inhibin B | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Insulin | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Leptin | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
LH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Methylmalonic acid | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Oestradiol | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Osteocalcin | Hormones in Serum | na | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
P3NP | Hormones in Serum | na | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
PIGF | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
PINP | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Progesteron | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Progesteron,17OH | Hormones in Serum | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Prolactin | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
PTH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Renin antigen | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
sFlt1 | Hormones in Serum | na | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
SHBG | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
T3 | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
T3, free | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
T4 | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
T4, free | Hormones in Serum | Quant. | Consensus***** | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
TBG | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Testosteron | Hormones in Serum | Quant. | Expert******* | Analytical by expert | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
TSH | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Vit. D 1,25 (OH)2 | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
Vitamin B12 | Hormones in Serum | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically different
3-MT | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human Urine. | Store at 2 - 8°C upon receiving.
5-HIAA | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Adrenaline | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Aldosteron | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Cortisol | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Cortison | Hormones in Urine | Qualit. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Dopamine | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Free metanefrine | Hormones in Urine | Qualit. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Free normetanephrine | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Free 3-MT | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
HVA | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Metanefrine | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Noradrenaline | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Normetanephrine | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
VMA | Hormones in Urine | na | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
AS aangetoond tegen bofvirus? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
AS aangetoond tegen SARS-CoV-2? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Beschermd tegen hepatitis A? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Beschermd tegen hepatitis B? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Beschermd tegen mazelen? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Beschermd tegen rubella? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Beschermd tegen waterpokken? | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Bof antibodies | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Bof IgG | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Hepatitis A IgG | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Hepatitis A IgT | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Hepatitis B Anti-HBs | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Measels antibodies | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Measels IgG | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Rubella antibodies | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Rubella IgG | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
SARS-CoV IgG | Immunity serology |  | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Varicella zoster antibodies | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Varicella zoster IgG | Immunity serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Ciclosporin | Immunosuppressant microsampling | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human EDTA blood | Store at -20°C upon receiving. (samples are shipped without specified shipping conditions) Samples are taken on the microsampling device supplied by yourself in accordance with the sampling method associated with the relevant microsampling device.
Creatinine | Immunosuppressant microsampling | Quant. | Expert | Analytical by expert (Sysmex measuring) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Everolimus | Immunosuppressant microsampling | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Hematocrit | Immunosuppressant microsampling | Quant. | Expert | Analytical by expert (Sysmex measuring) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycophenolic acid | Immunosuppressant microsampling | Quant. | Expert | Analytical by expert (LC/MS/MS method) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Sirolimus | Immunosuppressant microsampling | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Tacrolimus | Immunosuppressant microsampling | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CRP | Inflammation markers | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human Li-heparinized plasma and serum. | Store at 2 - 8°C upon receiving.
Procalcitonine | Inflammation markers | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Blastocystis spp. PCR | Intestinal protozoa Molecular Diagnostics | na | Expert | ALTM | Theoretically likely | Analytically different | Analytically different | Human feces. | Store at 4°C upon receiving. Analyse samples as quickly as possible after receipt.
Cryptosporidium spp. PCR | Intestinal protozoa Molecular Diagnostics | Qualit. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Dientamoeba fragilis PCR | Intestinal protozoa Molecular Diagnostics | Qualit. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Entamoeba dispar PCR | Intestinal protozoa Molecular Diagnostics | Qualit. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Entamoeba histolytica PCR | Intestinal protozoa Molecular Diagnostics | Qualit. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Giardia lamblia PCR | Intestinal protozoa Molecular Diagnostics | Qualit. | Expert | ALTM | Theoretically likely | Analytically different | Analytically different
Ectoparasite | Intestinal-, tissue and ecto parasites | Qualit. | Expert | Analytical by 6 experts | Theoretically likely | NA | NA | Digital image. | NA
Parasite intestinal/tissue | Intestinal-, tissue and ecto parasites | Qualit. | Expert | Analytical by 6 experts | Theoretically likely | Analytically different | Analytically different | Human feces and/or punctate | Store at room temperature upon receiving.
ALAT | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | The samples must be stored in the freezer (< -70°C) upon receiving. Once thawed, the samples are stable during 6 hours at 6-24°C. When stored at 2-8°C the samples are stable during 24 hours (in the light protective brown glass vial).
Alk. Phosphatase | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Amylase | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
ASAT | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Bilirubin | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Calcium | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Chloride | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Cholesterol | KILO | na | Consensus | ALTM | Theoretically likely | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CK | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Creatinine | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Gamma-GT | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Glucose | KILO | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Inorg. Phosphate | KILO | na | Consensus | Consensus - method group | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Iron | KILO | na | Consensus | Consensus - method group | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
LD | KILO | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Potassium | KILO | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Sodium | KILO | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Protein | KILO | Quant. | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Triglycerides | KILO | na | Consensus | ALTM | Theoretically likely | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urate | KILO | na | Reference | Reference method according JCTML list | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Urea | KILO | Quant. | Consensus | Consensus - method group | Analytically different | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Abberating population | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different | Human EDTA blood and smear | Store at room temperature upon receiving. Measure as soon as possible. (samples are shippped through courier)
B-lymphocytes | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 1 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 10 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 103 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 109 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 117 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 11a | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 11b | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 11c | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 123 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 13 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 133 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 135 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 138/Vs38c | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 14 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 15 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 16 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 18 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 19 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 2 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 20 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 200 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 21 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 22 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 23 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 235a | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 24 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 25 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 3 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 30 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 33 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 34 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 36 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 37 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 38 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 4 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 41 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 42 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 43 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 45 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 45RA | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 45RO | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 5 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 56 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 57 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 61 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 62P | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 64 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 65s | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 66 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 66e | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 7 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 71 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 79a | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 79b | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 8 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CD 9 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyCD 3 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyCD22 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyCD79a | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-A | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-D | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-G | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-Kappa | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-Lambda | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
CyIg-M | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
FMC-7 | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
HLADR | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
Monocytair | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
MPO | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
Myeloïd | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
NK-lymphocytes | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-A | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-D | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-G | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-Kappa | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-Lambda | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
sIg-M | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
TCR-alpha-beta | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
TCRgd | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
TdT | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
T-lymphocytes | Leukemia/Lymphoma typing | Quant. | Expert | Analytical by 5-6 experts (Averaging 5-6 experts lab, with no cons most common value) | Theoretically likely | Analytically different | Analytically different
Apo A1 lipoproteïn | Lipids | Quant. | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human serum. | Store at < -70°C upon receiving. (samples are shipped on dry-ice) Once thawed, the samples are stable for 6 hours at 16-24°C; stored at 2-8°C the samples are stable for 24 hours.
Apo B lipoproteïn | Lipids | Quant. | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Cholesterol | Lipids | Quant. | Reference | Reference method according JCTML list | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Cholesterol ratio | Lipids | Quant. | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
HDL-Cholesterol | Lipids | Quant. | Reference | Reference method according JCTML list | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
hsCRP | Lipids | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
LDL-Cholesterol | Lipids | Quant. | Expert | Analytical by expert (direct method) | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Lp(a) | Lipids | Quant. | Consensus | Consensus - method group | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
non-HDL cholesterol | Lipids | Quant. | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Triglycerides | Lipids | Quant. | Consensus | ALTM | Analytically different | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Albumine liquor | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human CSF and serum. | Store at -16 to -24°C upon receiving. (samples are shipped without specified shipping conditions)
Albumine serum | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG index | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG liquor | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG serum | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
intrathecale IgG fraction | Liquor protein | Quant. | Expert | Edited consensus by expert (median) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
intrathecale IgG fraction (Kwal.) | Liquor protein | Qualit. | Consensus | Edited consensus by expert (median) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
intrathecale IgG production | Liquor protein | Qualit. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Liquor bands | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Q Albumin | Liquor protein | Quant. and Qualit. | Expert | Edited consensus by expert (median) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Q IgG | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Serum bands | Liquor protein | Quant. | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Total protein liquor | Liquor protein | Quant. and Qualit. | Expert | Edited consensus by expert (median) | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
017 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
B. burgdorferi IgG aanwezig? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
B. burgdorferi IgG blot? | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
B. burgdorferi IgM aanwezig? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
B. burgdorferi IgM blot? | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Borrelia burgdorferi IgG | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Borrelia burgdorferi IgM | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Borrelia IgTotal | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CL1 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CL2 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CL3 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
DbpA (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
DbpA (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
iv 1 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
iv 2 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
iv 3 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
iv 4 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Lues infectie? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Lyme behandeling? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Non-treponemale antistoffen aanwezig? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Osp17 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
OspA (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
OspA (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
OspC (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
OspC (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p100 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p100 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p14 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p15 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p17 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p18 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p18 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p30 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p39 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p39 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p41 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p41 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p43 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p44.5 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p45 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p47 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p58 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p58 (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
p83 (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
T. pallidum IgG blot? | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
T. pallidum FTA. | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
T. pallidum IgM FTA | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
T. pallidum screening | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
T. pallidum TPHA/TPPA | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TmpA (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Tp15 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TP17 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Tp257 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Tp453 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TpN15 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TpN17 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TpN47 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
TPp47 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Treponemale antistoffen aanwezig? | Lues / Lyme serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
VDRL/RPR | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
vdrl1 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
vdrl2 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
vdrl3 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
vdrl4 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
vdrl5 (T. pal) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
VlsE (IgG blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
VLsE (IgM blot) | Lues / Lyme serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
% pos macrophagen | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different | Human BAL cells suspension and 2 cytospins. | Store at 2 - 8°C upon receiving.
CD103 | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD103,CD4/CD4 | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD19+ | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD3+ | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD3+CD4+ | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD3+CD8+ | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
CD3-CD16+CD56+ | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Eosinophils | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Golde score | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Incubation temperature | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Lymphocytes | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Lymphocytes on flow | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Macrophages | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Mast cells | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Neutrophils | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Number of cells | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Other | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Plasma cells | Lymphocyte subsets BAL | Quant. | Consensus | Edited consensus by expert | Analytically different | Analytically different | Analytically different
Alfa thal. - alfa 3.7 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human DNA | Store at 4°C upon receiving.
Alfa thal. - alfa 4.2 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --20.5 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --FIL | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --MED | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --SEA | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --THAI | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alfa thal. --Triplicatie | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CF-delta 508 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HFE C282Y | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HFE H63D | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HLA-B27 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HLA-DQ2.2 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HLA-DQ2.5 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
HLA-DQ8 | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
LCT -13910C>T | Molecular Biology various | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CALR Mutation | MPN molecular diagnostics | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different | Human gDNA. | Store at 4°C upon receiving.
CALR type of mutation | MPN molecular diagnostics | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different
CALR VAF | MPN molecular diagnostics | Quant. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different
JAK2 exon 12 mutation | MPN molecular diagnostics | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different
JAK2 V617F | MPN molecular diagnostics | Quant. and Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different
MPL mutation | MPN molecular diagnostics | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different
Albumin spectrum | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human serum. | Store as patientmaterial.
Alpha-1-globulin | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Alpha-2-globulin | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Beta-globulin | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Cryoglobulin | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Free kappa | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Free lambda | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Gamma-globulin | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA (light chain ?) | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA heavy chain | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA kappa | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA kappa 2nb band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA kappa 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA lambda | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA lambda 2nd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgA lambda 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgD | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgD (light chain ?) | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgD kappa | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgD lambda | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG (light chain ?) | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG heavy chain | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG kappa | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG kappa 2nb band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG kappa 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG lambda | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG lambda 2nd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgG lambda 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM (light chain ?) | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM heavy chain | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM kappa | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM kappa 2nd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM kappa 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM lamba 2nd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM lamba 3rd band | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
IgM lambda | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Ratio (kappa/lambda) | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
sVLK kappa | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
sVLK lambda | M-protein | Qualit. | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Total protein | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Total Protein (urin) | M-protein | na | Expert | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Determination | Mycobacteriology | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different | Culture medium. | Store at 2 - 8°C upon receiving. Treat sampels in BLS3 room.
Determination | Mycology | Qualit. | Expert | Analytical by 2 experts | Theoretically likely | Analytically different | Analytically different | Agar | Store at room temperature upon receiving.
Albumin | Neonatal Bilirubin | na | Consensus | Consensus - method group | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen human Serum. | Store at < -70°C upon receiving. (samples are shipped without specified shipping conditions). Protect from extreme light. Thawed samples are stable for 8 hours at 2-8°C.
Bilirubin | Neonatal Bilirubin | Quant. | Reference | Reference method according JCTML list | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Creatinine | Neonatal Bilirubin | na | Consensus | Consensus - method group | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Glucose (low) | Neonatal Bilirubin | Quant. | Consensus | Consensus - method group | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Hb in feces | Occult blood in feces | na | Consensus | Consensus - method group | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised feces suspension | Store at 2 - 8°C upon receiving. The shelf life of the reconstituted samples is 8 hours at 2-8°C.
Hb in feces (qual.) | Occult blood in feces | Qualit. | Expert | Analytical by expert | unknown | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Cabozantinib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
Desethylsunitinib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Imatinib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Olaparib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Pazopanib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sunitinib | Oral oncolytics | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
c.1236G>A | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Human gDNA. | Store at 4°C upon receiving.
c.1679T>G | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
c.2846A>T | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2C19*2 | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*3 | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*4 | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*41 | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*5 (deletie) | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*6 | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
CYP2D6*duplicatie | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
DPYD*2A (IVS14+1 G>A) | Pharmacogenetics | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Albumin | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Frozen Human serum. | Store at < -70°C upon receiving. (samples are shipped on dry-ice)
Alpha-1-antitrypsin | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Beta-2-microglobulin | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
C3 | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
C3d | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
C4 | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Ceruloplasmin | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
CRP | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Ferritin | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Haptoglobin | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgA | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgA1 | Plasma proteins | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgA2 | Plasma proteins | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgD | Plasma proteins | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgE | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgG | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgG1 | Plasma proteins | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgG2 | Plasma proteins | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgG3 | Plasma proteins | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgG4 | Plasma proteins | na | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
IgM | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Iron saturation % | Plasma proteins | Quant. | Consensus | ALTM | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Prealbumin | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Transferrin | Plasma proteins | Quant. | Reference | Reference method according JCTML list | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Zure alfa-1-glyco | Plasma proteins | na | Consensus | Consensus - method group | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
anti-GPIa/IIa | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different | Frozen human serum, trinatriumcitrate plasmaferese and K3EDTA blood. | The samples must be stored at -20°C upon receiving.(samples are shipped on dry-ice)
anti-GPIb/IX | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-GPIIb/IIIa | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-GPIV | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-15a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-15b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-1a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-1b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-2a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-2b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-3a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-3b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-4a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-4b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-5a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
anti-HPA-5b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
Aspecific | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HLA-Class-I antibodies | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-15a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-15b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-1a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-1b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-2a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-2b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-3a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-3b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-4a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-4b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-5a | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
HPA-5b | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
Platelet-antibodies (incl. HLA-ab) | Platelets / leukocyte serology | Qualit. | Expert | Analytical by 2 experts (standard platelet serology methods) | Theoretically likely | Analytically different | Analytically different
CD14 PNH clone (monocytes) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different | Human whole blood. | Store at room temperature (5 - 25°C) upon receiving.
CD157 PNH kloon (monocytes) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD157 PNH kloon (neutrophils) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD16 PNH kloon (neutrophils) | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD24 PNH kloon (neutrophils) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD55 PNH clone (monocytes) | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD55 PNH clone total | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD55 PNH kloon (neutrophils) | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD55 Type II cells | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD55 Type III cells | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD58 PNH clone total | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD58 Type II cells | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD58 Type III cells | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD59 PNH clone (monocytes) | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD59 PNH clone total | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD59 PNH kloon (neutrophils) | PNH diagnostics | na | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD59 Type II cells | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
CD59 Type III cells | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
FLAER PNH kloon (monocytes) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
FLAER PNH kloon (neutrophils) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
Monocytes | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
Monocytes (within WBC) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
Neutrophils | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
neutrophils (within WBC) | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
Red blood cells | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
White blood cells | PNH diagnostics | Quant. | Expert | Edited consensus by expert (median) | NA | Analytically different | Analytically different
3-OH-butyric acid | POCT in whole blood | na | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Analytically different | Human whole blood. | Store at 2 - 8°C upon receiving. Sample in ACU-DROP-II dripper, measure as soon as possible.
Creatinine | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
CRP. | POCT in whole blood | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
D-Dimer | POCT in whole blood | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Glucose. | POCT in whole blood | na | Consensus | Consensus - method group | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Hemoglobin. | POCT in whole blood | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
hs-Troponin I. | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
hs-Troponin T. | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Lactate | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
NT-proBNP. | POCT in whole blood | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Potassium | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Sodium | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Troponine T | POCT in whole blood | na | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Analytically different
Coproporphyrin I | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human urine. | Store at -16 to -24°C upon receiving.
Coproporphyrin I (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Coproprophyrin III | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Coproprophyrin III (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Creatinine | Porphyrins in urine | na | Expert | Analytical by 3 experts (Roche enzymatic) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
D-aminolevulinic acid | Porphyrins in urine | na | Expert | Analytical by 3 experts (Recipe) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
D-aminolevulinic acid (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (Recipe) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Heptaporphyrin | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Heptaporphyrin (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Hexaporphyrin | Porphyrins in urine | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Hexaporphyrin (per creatinine) | Porphyrins in urine | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Pentaporphyrin | Porphyrins in urine | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Pentaporphyrin (per creatinine) | Porphyrins in urine | na | Consensus | Consensus - method group | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Porphobilinogen | Porphyrins in urine | na | Expert | Analytical by 3 experts (Recipe) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Porphobilinogen (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (Recipe) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Coproporfyrin | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Coproporfyrin (per Creatinin) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Porphyrin | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Porphyrin (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Uroporphyrin | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Total Uroporphyrin (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Uroporphyrin I | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Uroporphyrin I (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Uroporphyrin III | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Uroporphyrin III (per creatinine) | Porphyrins in urine | na | Expert | Analytical by 3 experts (LC Fluorescence) | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Amitriptyline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Lyophilised calf serum. | Store at 2 - 8°C upon receiving.
Aripiprazol | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Citalopram | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Clomipramine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Clopentixol | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Clozapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Dehydro-aripiprazole | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desipramine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylcitalopram | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylclomipramine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylclozapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylfluoxetine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylmirtazapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylsertraline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Desmethylvenlafaxine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
E-10-OH-Nortriptyline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Fluoxetine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Fluvoxamine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Haloperidol | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Hydroxyrisperidon | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Imipramine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Lithium | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Mirtazapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Nortriptyline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Olanzapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Paroxetine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Quetiapine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Risperidon | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Sertraline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Venlafaxine | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
Z-10-OH-Nortriptyline | Psychotherapeutic drugs | Quant. | Reference | Reference - by weighing in a traceable standard in 0-serum. | Theoretically likely | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation
PT-INR | PT-INR for oral anticoagulation | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Analytically different | Lyophilised human citrate plasma. | Store at 2 - 8°C upon receiving.
Adenovirus antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. (4 pairs) | Store at 2 - 8°C upon receiving.
Adenovirus IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Adenovirus IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Adenovirus IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Chlamydia pneumoniae/psittacii antistoffen | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Chlamydia pneumoniae/psittacii IgA antistoffen | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Chlamydia pneumoniae/psittacii IgG antistoffen | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Chlamydia pneumoniae/psittacii IgM antistoffen | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii antibodies fase 1 | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii antibodies fase 2 | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii fase 1 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii fase 1 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii fase 2 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii fase 2 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Coxiella burnetii IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza A antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza A IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza A IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza A IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza B antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza B IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza B IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Influenza B IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Legionella IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Legionella IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Legionella IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Legionella IgTotaal antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycoplasma pneumoniae antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycoplasma pneumoniae IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycoplasma pneumoniae IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycoplasma pneumoniae IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Mycoplasma pneumoniae IgM antibodies (2e test) | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1 antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1 IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1-3 antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1-3 IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1-3 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 1-3 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 2 antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 2 IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 2 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 2 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 3 antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 3 IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 3 IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Parainfluenza 3 IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente adenovirus infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente Chlamydia infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente influenza A infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente influenza B infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente Legionella infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente Mycoplasma infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente parainfluenza infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente Q-koorts | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Past bij recente RSV infectie | Respiratory serology | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
RSV antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
RSV IgA antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
RSV IgG antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
RSV IgM antibodies | Respiratory serology | na | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Antibodies against CCP | Rheuma | Qualit. | Expert | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
antibodies against CCP ULN categorie | Rheuma | Qualit. | Expert | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Reumafactor ULN categorie | Rheuma | Qualit. | Expert | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Rheumatoid factor (IgM) | Rheuma | Qualit. | Expert | Edited consensus by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Acrosomale regio <= 2 vacuoles | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Acrosomale vacuoles < 20% kop-opp. | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Acrosoom 40-70% van kop-opp. | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Bevat geen knik | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Concentration | Semen analysis | Quant. | Consensus | Edited consensus by expert (Based on rounded consensus group) | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Conclusion | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Dikte vrijwel gelijk over de hele lengte | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
ERC < 1/3 van kop-opp. | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Geen post-acrosomale vacuoles | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Iets dunner dan het middenstuk | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Immotile | Semen analysis | na | Consensus | Edited consensus by expert (Based on rounded consensus group) | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Lengte as middenstuk = lengte as kop | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Lengte is circa 10x normale koplengte | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Morfologie (% normaal, WHO 2010) | Semen analysis | Quant. | Consensus | Edited consensus by expert (Based on rounded consensus group) | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Non-progressive | Semen analysis | Quant. | Consensus | Consensus - method group | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Ongeveer even lang als de kop | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Ovaal van vorm | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Progressive | Semen analysis | Quant. | Consensus | Consensus - method group | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Rapid progressive motile (%A) | Semen analysis | Quant. | Expert | Edited consensus by expert | Theoretically likely | Analytically different | Retrospectively based on between-lab variation | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Regelmatig van vorm | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Regelmatige, egale omtrek | Semen analysis | na | Expert | Theoretically by 2 experts | Theoretically likely | Analytically different | Analytically different | Digital image.
Total count | Semen analysis | na | Consensus | Consensus - method group | Theoretically likely | Analytically different | Analytically different | Human pool sperm sample/smears, sperm sample/video | Store at 4 - 8°C upon receiving. Samples are stable for week.
Soluble IL2 receptor | Soluble IL2 receptor | Quant. | Consensus | ALTM | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Store at -20 to -80°C upon receiving.(samples are shipped on dry-ice)
Chlamydia trachomatis (real-time PCR) | STD External Molecular Quality Assessment Scheme | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different | Urogenital swab in stabilisation solution. | Store at 4°C upon receiving. Measure as soon as possible.
Neisseria gonorrhoeae (real-time PCR) | STD External Molecular Quality Assessment Scheme | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
Trichomonas vaginalis (real-time PCR) | STD External Molecular Quality Assessment Scheme | Qualit. | Expert | Analytical by expert | Theoretically likely | Analytically different | Analytically different
antibodies against Intrinsic factor | Stomach / Liver antibodies | Qualit. | Expert | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Store at -16 to -24°C upon receiving.(samples are shipped without specified shipping conditions)
antibodies against mitochondria | Stomach / Liver antibodies | Qualit. | Expert | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against parietal cell | Stomach / Liver antibodies | Qualit. | Expert | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against smooth muscle | Stomach / Liver antibodies | Qualit. | Expert | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
6-MMP | Thiopurines | Quant. | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different | Frozen human pooled serum | Store at < -20°C upon receiving. (samples are shipped on dry-ice)
6-TGN | Thiopurines | Quant. | Consensus | ALTM | Theoretically likely | Analytically different | Analytically different
Factor II 20210 | Thrombophilia (genetics) | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Humaan gDNA. | Store at 4°C upon receiving.
Factor V Leiden | Thrombophilia (genetics) | Qualit. | Expert | Analytical by expert | Theoretically likely | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against Tg | Thyroid Antibodies | Qualit. | Consensus | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation | Frozen human serum. | Store at -16 to -24°C upon receiving. (samples are shipped without specified shipping conditions)
antibodies against TPO | Thyroid Antibodies | Qualit. | Consensus | Edited consensus by expert | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
antibodies against TSH-R | Thyroid Antibodies | Qualit. | Consensus | Consensus - method group | unknown | Retrospectively based on between-lab variation | Retrospectively based on between-lab variation
Toxic substance | Toxicology | Qualit. | Reference | Reference - by weighing in a traceable standard in 0-serum. |  | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons |  | Lyophilised calf serum. | Store at 2 - 8°C upon receiving.
Tryptase | Tryptase | Quant. | Consensus | Consensus - method group | unknown | Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Retrospectively based on between-lab variation | Frozen human pooled serum diluted in omniplasma | Sore at -16 to -24°C upon receiving.(samples are shipped without specified shipping conditions)
3-Metoxytyramine | Tumor markers | Quant. | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different | Frozen human serum. | Store at < -70°C upon receiving.(samples are shipped on dry-ice)
AFP | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
CA-125 | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
CA-15.3 | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
CA-19.9 | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Calcitonin | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
CEA | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Chromogranin A | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Gastrin | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
HCG plus HCG-beta | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
HE4 | Tumor markers | na | Consensus | ALTM | Analytically different | Analytically different | Analytically different
Metanefrine | Tumor markers | Quant. | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Normetanefrine | Tumor markers | Quant. | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
PSA total | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
PSA, free | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
PSA, ratio free/total | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Serotonin | Tumor markers | na | Consensus | ALTM | Analytically different | Analytically different | Analytically different
Tg-Autoantibodies | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Thyreoglobulin | Tumor markers | na | Consensus | Consensus - method group | Analytically different | Analytically different | Analytically different
Urine element | Urine Sediment (qualtative) | Qualit. | Expert | Analytical by expert | NA | NA | NA | Digital image. | NA
Beta-carotene | Vitamin A/E and beta-carotene | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human serum. | Store at 2 - 8°C upon receiving.
Vitamin A | Vitamin A/E and beta-carotene | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Vitamin E | Vitamin A/E and beta-carotene | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
Pyridoxalphosphate (PLP) | Vitamin B/B6 | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | Lyophilised human blood | Store at 2 - 8°C upon receiving.
TPP (Th. Diphosphate) | Vitamin B/B6 | Quant. | Consensus | ALTM | Theoretically likely | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons | analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
*We do not distinguish between pooled and single donor samples; only when relevant, it will be indicated.
Also retrospectively based on between-lab variation

## Wijzgingen
Datum | Initialen | Omschrijving
20191127 | RM | Namen rondzendingen en bepalingen aangepast, α wordt Alfa, ß wordt Bèta. Aantal eenheden naar UCUM-standaard.
20200210 | RM | Nieuwe bepalingen hormonen in serum, Hormonen in urine en tumormerkstoffen toegevoegd. TE voor HbA1C en Hemocytomtrie aangepast op basis EFML database.
20200227 | RM | bijgewerkt met nieuwe TE op basis van EFML database
20200416 | RM | Pathologie rondzendingen verwijderd
20200430 | RM | Liquor eiwitten aangepast
20202705 | RM | Bepalingen cholesterol ratio en non-hdl cholesterol toegevoegd
20200703 | RM | Urine kwalitatief alle bepalingen behalve kreatinine, (micro)albumine, Amylase en oxalaat op Expert waarde ipv consensus
20200703 | RM | Homocysteine aangepast van Consensus, ALTM naar Reference, weighted
20200911 | RM | intercellulair (ICS) aangepast naar Extracellulair oppervlak keratinocyten (ECS) bij de rondzending antistoffen tegen huid
20200918 | RM | rondzending Immunosuppressant microsampling toegevoegd.
20201103 | RM | Eenheden in UCUM notatie gezet.
20201117 | RM | Nieuwe gevoeligheid Oxacilline bacteriologie toegevoegd, rondzending namen hernoemd volgens mail van 16112020 Guid van Dam
20201123 | RM | Eenheden sectie endocrinlogie aangepast.
20201130 | RM | Eenheden, nieuwe rondzendingen en niuewe pilot rondzendingen toegevoegd. Alle eenheden en Tea en Sdsa gecontroleerd em aangepast.
20201209 | RM | Eenheid C4 bij complement factoren en plasma eiwitten aangepast van mg/L naar g/L
20201218 | RM | Kreatinine bepaling bij rondzending Immunosuppressiva microsampling toegevoegd en de bepalingen ASXL1 mutatie, RUNX1 mutatie en TP53 mutatie bij rondzending AML molecular diagnostics toegevoegd.
20201224 | RM | Pilot blood pigments in cerebrospinal fluid hernoemd naar Blood pigments in cerebrospinla fluid
20201229 | RM | Alle Target value, Source target value en commutabillity kolomen gecontroleerd en bijgewerkt.
20210129 | RM | Bloed/darm parasieten spliting verwerkt. Nu twee rondzendingen bloedparasieten en Darm-, weefsel- en ectoparasieten.
20210204 | RM | Nieuwe bepalingen Benzylpenicilline en Hydroxychloroquine bij Geneesmiddelen afwisselend toegevoegd.
20210223 | RM | Eenheid uraat aangepast rondzendingen Kilo en Klinische Chemie, bloed
20210224 | RM | Kolom source target value bijgewerkt.
20210316 | RM | Kolom "score based on" verwijderd. Kolom "Base for Quantitative score" en kolom "scored items" bijgemaakt en bijgewerkt.
20210326 | RM | TE bijgewerkt met laatst biologische variatie volgens de ELFM database van jan 2021
20210329 | RM | Kolom "Base for Quantitative score" en kolom "scored items" bijgewerkt.
20210412 | RM | Kolom "LAST TE CHANGE" aangemaakt en ingevuld.
20210622 | RM | Bepaling protein-unbound (free) flucloxacillin Antimicrobiële middelen 2 toegevoegd
20210715 | RM | Placeholder "Reuma" die een lijst van veel verschillende eenheden toonde vervangen door [IU]/ml. IVM overgang naar UCUM notatie.
20210715 | RM | -target value hernoemd naar target value type -source target value hernoemd naar metrological traceability target value -Kolommen stability, homogenity, material type, material extra info toegevoegd
20210715 | RM | TEa & SA tabblad en sample aspects tabblad gemaakt met resp TEa/SA waarden en monster aspecten
20210720 | RM | Bij AML molecular diagnostics bepaling AML-ETO1 hernoemd naar RUNX1-RUNX1T1
20210824 | RM | kolommen stability, homogeniteit, material type en material extra info bijgewerkt.
20211209 | RM | Vaginaal fibronectine aangepast naar Fetal fibronectin
20211215 | RM | nieuwe bepalingen hormonen in speeksel en hormone in urine toegevoegd.
20211216 | RM | Bepaling lactaat poct in volbloed toegevoegd. Nieuwe rondzending Urine sediment toegevoegd.
20211217 | RM | Bepaling hsCRP lipiden en Gecorrigeerd calcium klinische chemie bloed toegevoegd.
20211221 | RM | Rondzending pilot Acanthamoeba Extern. Mol. Quality Assessment hernoemd naar Acanthamoeba Extern. Mol. Quality Assessment
20220119 | RM | Bepaling Estrone toegevoegd hormonen in serum.
20220120 | RM | Pilot bilirubine in volbloed hernoemd naar Bilirubine in volbloed.
20220131 | RM | CDT, busulfan SA bijgewerkt.
20220208 | RM | Chimerism SA bijgewerkt.
20220208 | RM | E-10-OH-Nortriptyline en Desmethylsertraline bij psychofarmac verwijderd. Bepaling %a bij semen toegevoegd. Sunitinib toegvoegd bij orale oncolytica, kalfsserum aangepast naar humaan serum. Ritonavir verwijderd bij Anti-HIV middelen.
20220209 | RM | ACE, alcoholen/GHB, Anti-HIV middelen, Antituberculose middelen, Antivirale middelen, biologicals, bloedpigmenten liquor,busulfan, cardiaca, CDT, Hb varianten SA bijgewerkt.
20220210 | RM | HbA1C volbloed, Hormonen in serum SA en TE aangepast. AF-isoenzymen, Hormonen in speeksel, Hormonen urine SA aangepast.
20220211 | RM | ANCA-GBM, Antischimmel middelen, APL molecular diagnostics, benzodiazepinen, Anti-epileptica vrije concentratie, Anti-epileptica SA aangepast. Vitamine A/E en bèta-caroteen SA en TE aangepast
20220212 | RM | Vitamine B1/B6, Bilirubine in volbloed, Fructosamine SA aangepast. Bloedgassen SA en TEa aangepast.
20220212 | RM | Ethanol in serum, Feces-chemie, Foetaal fibronectine, Galzure zouten, hartmerkers SA aangepast.
20220215 | RM | Foetomaternale transfusie (FMT), hemoximetry, homocysteine SA aangepast
20220216 | RM | Klinische Chemie, urine kwantitatief, lipiden, maag/lever antistoffen SA aangepast
20220217 | RM | Tumormerkstoffen SA aangepast
20220218 | RM | Antimicrobiële middelen, Coeliakie, Moleculaire diagnostiek bij MPN, M-proteïne, occult bloed in feces, onstekingsmerkers SA aangepast.
20220218 | RM | Drugs of Abuse toegevoegd en beplaing Teicoplanine geneesmiddelen afwisslend toegevoegd
20220221 | RM | POCT in volbloed SA en TE aangepast. Porfyrines, thiopurines en psychofarmaca SA aangepast.
20220222 | RM | PT-INR voor orale antistolling, soluble Il2 receptor, semen, darmprotozoa moleculaire diagnostiek en stolling SA bijgewerkt.
20220307 | RM | Pijnstillers, complement factoren, diabetes en SOA moleculaire diagnostiek SA bijgewerkt
20220307 | RM | Plasma eiwitten, liquor eiwitten en neonatale bilirubine SA en TE bijgewerkt.
20220308 | RM | Klinische chemie, bloed SA en TE bijgewerkt.
20220308 | RM | Antimicrobiële middelen 2 SA bijgewerkt.
20220406 | RM | Diabetes PP bijgewerkt
20220414 | RM | Tabblad sample aspects DOA bijgewerkt.
20220420 | RM | Rondzending SARS-CoV-2 toegevoegd.
20220428 | RM | Rondzending Tryptase en Allergy type III antistoffen toegevoegd
20220503 | RM | Rondzending semen bepaling Rapid progressive motile (%A) toegevoegd
20220511 | RM | Bepaling SARS-CoV IgTotaal bij rondzending SARS-CoV-2 serologie toegevoegd
20220511 | RM | Drugs of abuse bijgewerkt.
20220610 | RM | Bepaling 3-MT hormonen in urine toegevoegd.
20220627 | RM | Merker CD 200 toegevoegd aan de rondzending leukemie/lymfoom typering
20220719 | RM | Toxicologie toegevoegd
20220729 | RM | LDL-cholesterol voortaan op expert gebaseerd op direct methode
20220808 | RM | Drugs of abuse target value type en metrological traceability value aangepast
20220912 | RM | Kolom material type bijgewerkt
20220915 | RM | Bepaling 11 deoxycorticosterone hormonen in serum toegevoegd
20220919 | RM | Tryptase en allergy type III bijgewerkt.
20221017 | RM | Baclofen toegevoegd.
20221020 | RM | Bepaling Mycoplasma pneumoniae IgM antistoffen (2e test) toegvoegd bij respiratoire serologie
20221102 | RM | Homogeniteit geneesmiddelenanalyse bijgewerkt.
20221122 | RM | Rondzending naam porfyrines aangepast naar porfyrines in urine.
20221207 | RM | kolommen stability, homogeniteit, material type en material extra info bijgewerkt.
20221208 | RM | Chromogranin A toegevoegd bij tumormerkstoffen
20221213 | RM | Kolom commutability bijgewerkt
20230112 | RM | Rondzendng BCR-ABL verwijderd.
20230113 | RM | anti-GPV bij Trombocyten/leukocyten serologie verwijderd wordt niet meer gedaan
20230116 | RM | Biologische variatie RET-He bijgewerkt
20230201 | RM | Desmethylsertraline toegevoegd bij de rondzending psychofarmaca
20230209 | RM | Biologische variatie bijgewerkt.
20230213 | RM | Nieuwe rondzending Bartonella henselae serologie toegevoegd.
20230213 | RM | Bepaling SARS-CoV IgG toegevoegd aan de rondzending immuniteit serologie.
20230214 | RM | Biologische variatie bijgewerkt.
20230215 | RM | Vedolizumab toegevoegd bij biologicals.
20230215 | RM | porfyrines in urine bepalingen porfyrine per kreatinie toegevoegd.
20230216 | RM | Orale oncolytica bijgewerkt.
20230221 | RM | Bilirubine neonatale bilirubine basis TEa aangepast naar Clinical decision
20230222 | RM | Albumine Klinische Chemie, bloed: Target value type en metrological traceability target value aangepast
20230223 | RM | Biologische variatie tumormerkstoffen bijgewerkt
20230301 | RM | Lp(a) lipiden aangepast naar nieuwe eenheid nmol/L
20230301 | RM | Lijst gevoeligheden bacteriologie bijgewerkt
20230324 | RM | Bepaling CALR VAF bij Moleculaire diagnostiek bij MPN toegevoegd.
20230405 | RM | Virologie rondzendingen (Virusdetectie in blaasjes/biopt/ooguitstrijk, Virusdetectie in feces, Virusdetectie in liquor en Virusdetectie in respiratoire materialen) verwijderd.
20230526 | RM | troponine I hartmerkers verwijderd.
20230606 | RM | monstermateriaal occult bloed feces aangepast
20230613 | RM | Busulfan monstermateriaal aangepast.
20230717 | RM | Materiaal thiopurines bijgewerkt, gevoeligheden bacteriologie bijgewerkt.
20230818 | RM | Nieuwe bepalingen CEBPA bZIP mutatie, BCOR mutatie, EZH2 mutatie, GATA2 mutatie, IDH1 mutatie, IDH2 mutatie, KIT mutatie, SF3B1 mutatie, SRSF2 mutatie, STAG2 mutatie, U2AF1 mutatie, ZRSR2 mutatie bij AML-molecular diagnostics toegevoegd
20230408 | RM | PNH diagnostiek kolom metrologische traceerbaarheid doel waarde bijgewerkt.
20230818 | RM | Homogeniteit POCT in volbloed bijgewerkt. Was Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons. Is geworden Retrospectively based on between-lab variation
20230914 | RM | Bacteriologie lijst gevoeligheden en resistentiemechanisme bijgewerkt
20230918 | RM | Bepaling brucella serologie verwijderd, wordt niet meer uitgevoerd.
20230926 | RM | kolommen MATERIAL TYPE en MATERIAL EXTRA INFO bijgewerkt.
20231011 | RM | Tryptase homogeniteit aangepast naar Analytically according CLSI-EP30A Characterization and Qualification of Commutable Reference Materials for Laboratory Medicine; Approved Guideline; of ISO 3528 Statistical methods for use in proficiency testing by interlaboratory comparisons
20231025 | RM | Beenmergmorfologie nieuwe opzet verwerkt
20231116 | RM | Bepalingen Cabozantinib en Olaparib toegvoegd bij Orale Oncolytica. Bij psychofarmaca Desmethylclozapine, Dehydro-aripiprazole en E-10-OH-nortriptyline
20231127 | RM | Bepaling TG autoantibodies kwal verwijderd bij tumormerkstoffen
20231207 | RM | Bacteriologie lijst gevoeligheden en resistentiemechanisme bijgewerkt
20231112 | RM | Bepaling 21-deoxycortisol toegevoegd aan de rondzending hormone in speeksel
20231219 | RM | Stolling bepaling afgeleide fibronogeen verwijderd, fibrinogeen vlgs Clauss hernoemd naar fibrinogeen
20231221 | RM | Totaal Coproporfyrine en Totaal Coproporfyrine (per kreatinine) toegevoegd bij rondzending Porfyrines in urine
20231221 | RM | Rondzending APL molecular diagnostics verwijderd gaat voortan via UKNEQAS
20240115 | RM | Sample aspects porfyrines, kolommen TARGET VALUE TYPE en METROLOGICAL TRACEABILITY TARGET VALUE bijgewerkt
20240216 | RM | Bloed cel morfology bijgewerkt
20240219 | RM | Beplaing Cefuroxime toegevoed bij rondzending Geneesmiddelen afwisselend
20240301 | RM | Rondzending bloedcelmorfologie bijgewerkt.
20240306 | RM | Teksten over het bepalen van de expertwaarden hormonen in serum bijgewerkt.